## Phase 4
Phase 4 of the calibration trip has two traveling systems. Updated 10 July 2015
All RAWDIF values computed with dclrinex version 17/03/2015. All P1-P2 values expressed with 2 digits for better comparison of 2.545*P1-1.545*P2 with P3/PPP.
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [5 empty] | Reference delays measured in the actual set-up. | [8 empty] | Measured / known antenna cable delays
[8 empty] | Updated 10 July 2015 | [5 empty] | The column "Note" is present to match the Table in the report | [8 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [6 empty] | The colums below make up Table 4 (Traveling - Visited) | [8 empty] | The colums below make up Table 5 (Traveling - Reference) | [8 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
[30 empty] | BP0R reference values TM243 |  | 221.5 | 224.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | CABDLYV | CABDLYR | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | L1 (ns) | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | TOTDLY(P3)
| 57046 | 57050 | BP1C | NIST | NIST | 57046-57050 | BP1C-NIST | -779.24 | 0.1 | -773.92 | 0.1 | BP1C-NIST | 57046-57050 | 939.1 | 80 | * | -779.24 | 79.86 | -773.92 | 85.18 | NIST-BP0R via BP1C | 2015.0685 | 275.5 | 133.4 |  | -151.53 | -293.63 | -153.5 | -295.6 | NIST wrt BP0R via BP1C | 2015.0685 | -72.13 | -71.1 | 121.7776
| 57046 | 57050 | BP0U | NIST | NIST | 57046-57050 | BP0U-NIST | -821.78 | 0.1 | -816.32 | 0.1 | BP0U-NIST | 57046-57050 | 743.4 | 80 | * | -821.78 | -158.38 | -816.32 | -152.92 | NIST-BP0R via BP0U | 2015.0685 | 275.5 | 133.4 |  | -152.48 | -294.58 | -154.48 | -296.58 | NIST wrt BP0R via BP0U | 2015.0685 | -73.08 | -72.08 | 120.874
| 57046 | 57054 | BP1C | NIS3 | NIST | 57046-57054 | BP1C-NIS3 | 599.95 | 0.2 | 617.6 | 0.2 | BP1C-NIS3 | 57046-57054 | 939.1 | 1545.8 | * | 599.95 | -6.75 | 617.6 | 10.9 | NIS3-BP0R via BP1C | 2015.074 | 298.5 | 133.4 |  | -64.92 | -230.02 | -79.22 | -244.32 | NIS3 wrt BP0R via BP1C | 2015.074 | -8.52 | -19.82 | -1238.3502
| 57046 | 57054 | BP0U | NIS3 | NIST | 57046-57054 | BP0U-NIS3 | 557.58 | 0.2 | 575.29 | 0.1 | BP0U-NIS3 | 57046-57054 | 743.4 | 1545.8 | * | 557.58 | -244.82 | 575.29 | -227.11 | NIS3-BP0R via BP0U | 2015.074 | 298.5 | 133.4 |  | -66.04 | -231.14 | -80.29 | -245.39 | NIS3 wrt BP0R via BP0U | 2015.074 | -9.64 | -20.89 | -1239.5475
| 57046 | 57054 | BP1C | NIS4 | NIST | 57046-57054 | BP1C-NIS4 | 572.35 | 0.2 | 589.56 | 0.2 | BP1C-NIS4 | 57046-57054 | 939.1 | 1516.5 | * | 572.35 | -5.05 | 589.56 | 12.16 | NIS4-BP0R via BP1C | 2015.074 | 298 | 133.4 |  | -66.62 | -231.22 | -80.48 | -245.08 | NIS4 wrt BP0R via BP1C | 2015.074 | -9.72 | -20.58 | -1211.4304
| 57046 | 57054 | BP0U | NIS4 | NIST | 57046-57054 | BP0U-NIS4 | 529.98 | 0.2 | 547.18 | 0.1 | BP0U-NIS4 | 57046-57054 | 743.4 | 1516.5 | * | 529.98 | -243.12 | 547.18 | -225.92 | NIS4-BP0R via BP0U | 2015.074 | 298 | 133.4 |  | -67.74 | -232.34 | -81.48 | -246.08 | NIS4 wrt BP0R via BP0U | 2015.074 | -10.84 | -21.58 | -1212.736
Note: Check C1-P1 option, if bias has been applied
First set-up at USNO, not to be used.
| 57067 | 57076 | BP1C | USN6 | USNO | 57067-57076 | BP1C-USN6 | -31.3 | 0.1 | -22.9 | 0.2 | BP1C-USN6 | 57067-57076 | 322.1 | 0 | ** | -31.3 | 290.8 | -22.9 | 299.2 | USN6-BP0R via BP1C | 2015.1329 | 0 | 133.4 | * | -362.47 | -229.07 | -367.52 | -234.12 | USN6 wrt BP0R via BP1C | 2015.1329 | -7.57 | -9.62 | -4.4007
| 57067 | 57076 | BP0U | USN6 | USNO | 57067-57076 | BP0U-USN6 | -66.7 | 0.1 | -58.2 | 0.1 | BP0U-USN6 | 57067-57076 | 118.8 | 0 | ** | -66.7 | 52.1 | -58.2 | 60.6 | USN6-BP0R via BP0U | 2015.1329 | 0 | 133.4 | * | -362.96 | -229.56 | -368 | -234.6 | USN6 wrt BP0R via BP0U | 2015.1329 | -8.06 | -10.1 | -4.9062
| 57067 | 57076 | BP1C | USN7 | USNO | 57067-57076 | BP1C-USN7 | -31.7 | 0.1 | -23.3 | 0.2 | BP1C-USN7 | 57067-57076 | 322.1 | 0 | ** | -31.7 | 290.4 | -23.3 | 298.8 | USN7-BP0R via BP1C | 2015.1329 | 0 | 133.4 | * | -362.07 | -228.67 | -367.12 | -233.72 | USN7 wrt BP0R via BP1C | 2015.1329 | -7.17 | -9.22 | -4.0007
| 57067 | 57076 | BP0U | USN7 | USNO | 57067-57076 | BP0U-USN7 | -67.1 | 0.1 | -58.6 | 0.1 | BP0U-USN7 | 57067-57076 | 118.8 | 0 | ** | -67.1 | 51.7 | -58.6 | 60.2 | USN7-BP0R via BP0U | 2015.1329 | 0 | 133.4 | * | -362.56 | -229.16 | -367.6 | -234.2 | USN7 wrt BP0R via BP0U | 2015.1329 | -7.66 | -9.7 | -4.5062
Second set-up at USNO, to be used.
| 57091 | 57097 | BP1C | USN6 | USNO | 57091-57097 | BP1C-USN6 | -31.54 | 0.1 | -22.97 | 0.1 | BP1C-USN6 | 57091-57097 | 322.1 | 0 | ** | -31.54 | 290.56 | -22.97 | 299.13 | USN6-BP0R via BP1C | 2015.1945 | 0 | 133.4 | * | -362.23 | -228.83 | -367.45 | -234.05 | USN6 wrt BP0R via BP1C | 2015.1945 | -7.33 | -9.55 | -3.8979
| 57091 | 57097 | BP0U | USN6 | USNO | 57091-57097 | BP0U-USN6 | -67.18 | 0.1 | -58.41 | 0.1 | BP0U-USN6 | 57091-57097 | 118.8 | 0 | ** | -67.18 | 51.62 | -58.41 | 60.39 | USN6-BP0R via BP0U | 2015.1945 | 0 | 133.4 | * | -362.48 | -229.08 | -367.79 | -234.39 | USN6 wrt BP0R via BP0U | 2015.1945 | -7.58 | -9.89 | -4.0087
| 57091 | 57097 | BP1C | USN7 | USNO | 57091-57097 | BP1C-USN7 | -31.92 | 0.1 | -23.41 | 0.1 | BP1C-USN7 | 57091-57097 | 322.1 | 0 | ** | -31.92 | 290.18 | -23.41 | 298.69 | USN7-BP0R via BP1C | 2015.1945 | 0 | 133.4 | * | -361.85 | -228.45 | -367.01 | -233.61 | USN7 wrt BP0R via BP1C | 2015.1945 | -6.95 | -9.11 | -3.6106
| 57091 | 57097 | BP0U | USN7 | USNO | 57091-57097 | BP0U-USN7 | -67.6 | 0.1 | -58.88 | 0.1 | BP0U-USN7 | 57091-57097 | 118.8 | 0 | ** | -67.6 | 51.2 | -58.88 | 59.92 | USN7-BP0R via BP0U | 2015.1945 | 0 | 133.4 | * | -362.06 | -228.66 | -367.32 | -233.92 | USN7 wrt BP0R via BP0U | 2015.1945 | -7.16 | -9.42 | -3.666
First set-up at OP, same reference as OPMT
| 57140 | 57147 | BP1C | OPMT/1 | OP | 57140-57147 | BP1C-OPMT/1 | -304.58 | 0.1 | -310.67 | 0.1 | BP1C-OPMT/1 | 57140-57147 | 222.8 | 100.1 |  | -304.58 | -181.88 | -310.67 | -187.97 | OPMT/1-BP0R via BP1C | 2015.3301 | 156.5 | 133.4 |  | 110.21 | 87.11 | 119.65 | 96.55 | OPMT/1 wrt BP0R via BP1C | 2015.3301 | 308.61 | 321.05 | 345.7778
| 57140 | 57147 | BP0U | OPMT/1 | OP | 57140-57147 | BP0U-OPMT/1 | -346.6 | 0.1 | -352.47 | 0.1 | BP0U-OPMT/1 | 57140-57147 | 26.1 | 100.1 |  | -346.6 | -420.6 | -352.47 | -426.47 | OPMT/1-BP0R via BP0U | 2015.3301 | 156.5 | 133.4 |  | 109.74 | 86.64 | 119.07 | 95.97 | OPMT/1 wrt BP0R via BP0U | 2015.3301 | 308.14 | 320.47 | 345.4778
| 57140 | 57147 | BP1C | OPM7/1 | OP | 57140-57147 | BP1C-OPM7/1 | -80.61 | 0.1 | -78.5 | 0.1 | BP1C-OPM7/1 | 57140-57147 | 222.8 | 128.1 |  | -80.61 | 14.09 | -78.5 | 16.2 | OPM7/1-BP0R via BP1C | 2015.3301 | 0 | 133.4 | * | -85.76 | 47.64 | -84.52 | 48.88 | OPM7/1 wrt BP0R via BP1C | 2015.3301 | 269.14 | 273.38 | 134.485
| 57140 | 57147 | BP0U | OPM7/1 | OP | 57140-57147 | BP0U-OPM7/1 | -122.65 | 0.1 | -120.28 | 0.1 | BP0U-OPM7/1 | 57140-57147 | 26.1 | 128.1 |  | -122.65 | -224.65 | -120.28 | -222.28 | OPM7/1-BP0R via BP0U | 2015.3301 | 0 | 133.4 | * | -86.21 | 47.19 | -85.12 | 48.28 | OPM7/1 wrt BP0R via BP0U | 2015.3301 | 268.69 | 272.78 | 134.2669
| 57140 | 57147 | BP1C | OPM8/1 | OP | 57140-57147 | BP1C-OPM8/1 | -84.16 | 0.1 | -82.07 | 0.1 | BP1C-OPM8/1 | 57140-57147 | 222.8 | 124.6 |  | -84.16 | 14.04 | -82.07 | 16.13 | OPM8/1-BP0R via BP1C | 2015.3301 | 0 | 133.4 | * | -85.71 | 47.69 | -84.45 | 48.95 | OPM8/1 wrt BP0R via BP1C | 2015.3301 | 269.19 | 273.45 | 138.004
| 57140 | 57147 | BP0U | OPM8/1 | OP | 57140-57147 | BP0U-OPM8/1 | -126.2 | 0.1 | -123.84 | 0.1 | BP0U-OPM8/1 | 57140-57147 | 26.1 | 124.6 |  | -126.2 | -224.7 | -123.84 | -222.34 | OPM8/1-BP0R via BP0U | 2015.3301 | 0 | 133.4 | * | -86.16 | 47.24 | -85.06 | 48.34 | OPM8/1 wrt BP0R via BP0U | 2015.3301 | 268.74 | 272.84 | 137.8014
Second set-up at OP, reference UTC(OP)
| 57149 | 57154 | BP1C | OPMT/2 | OP | 57149-57154 | BP1C-OPMT/2 | -601.62 | 0.1 | -607.63 | 0.1 | BP1C-OPMT/2 | 57149-57154 | 518.9 | 100.1 |  | -601.62 | -182.82 | -607.63 | -188.83 | OPMT/2-BP0R via BP1C | 2015.3521 | 156.5 | 133.4 |  | 111.15 | 88.05 | 120.51 | 97.41 | OPMT/2 wrt BP0R via BP1C | 2015.3521 | 309.55 | 321.91 | 346.8414
| 57149 | 57154 | BP0U | OPMT/2 | OP | 57149-57154 | BP0U-OPMT/2 | -647.12 | 0.1 | -652.81 | 0.1 | BP0U-OPMT/2 | 57149-57154 | 325.6 | 100.1 |  | -647.12 | -421.62 | -652.81 | -427.31 | OPMT/2-BP0R via BP0U | 2015.3521 | 156.5 | 133.4 |  | 110.76 | 87.66 | 119.91 | 96.81 | OPMT/2 wrt BP0R via BP0U | 2015.3521 | 309.16 | 321.31 | 346.7761
| 57149 | 57154 | BP1C | OPM7/2 | OP | 57149-57154 | BP1C-OPM7/2 | -377.66 | 0.1 | -375.52 | 0.1 | BP1C-OPM7/2 | 57149-57154 | 518.9 | 128.1 |  | -377.66 | 13.14 | -375.52 | 15.28 | OPM7/2-BP0R via BP1C | 2015.3521 | 0 | 133.4 | * | -84.81 | 48.59 | -83.6 | 49.8 | OPM7/2 wrt BP0R via BP1C | 2015.3521 | 270.09 | 274.3 | 135.4813
| 57149 | 57154 | BP0U | OPM7/2 | OP | 57149-57154 | BP0U-OPM7/2 | -423.1 | 0.1 | -420.58 | 0.1 | BP0U-OPM7/2 | 57149-57154 | 325.6 | 128.1 |  | -423.1 | -225.6 | -420.58 | -223.08 | OPM7/2-BP0R via BP0U | 2015.3521 | 0 | 133.4 | * | -85.26 | 48.14 | -84.32 | 49.08 | OPM7/2 wrt BP0R via BP0U | 2015.3521 | 269.64 | 273.58 | 135.4488
| 57149 | 57154 | BP1C | OPM8/2 | OP | 57149-57154 | BP1C-OPM8/2 | -381.23 | 0.1 | -379.1 | 0.1 | BP1C-OPM8/2 | 57149-57154 | 518.9 | 124.6 |  | -381.23 | 13.07 | -379.1 | 15.2 | OPM8/2-BP0R via BP1C | 2015.3521 | 0 | 133.4 | * | -84.74 | 48.66 | -83.52 | 49.88 | OPM8/2 wrt BP0R via BP1C | 2015.3521 | 270.16 | 274.38 | 139.0359
| 57149 | 57154 | BP0U | OPM8/2 | OP | 57149-57154 | BP0U-OPM8/2 | -426.67 | 0.1 | -424.16 | 0.1 | BP0U-OPM8/2 | 57149-57154 | 325.6 | 124.6 |  | -426.67 | -225.67 | -424.16 | -223.16 | OPM8/2-BP0R via BP0U | 2015.3521 | 0 | 133.4 | * | -85.19 | 48.21 | -84.24 | 49.16 | OPM8/2 wrt BP0R via BP0U | 2015.3521 | 269.71 | 273.66 | 139.0033
| 57181 | 57187 | BP1C | PTBB | PTB | 57181-57187 | BP1C-PTBB | -528 | 0.1 | -537.83 | 0.1 | BP1C-PTBB | 57181-57187 | 280.6 | 74 |  | -528 | -321.4 | -537.83 | -331.23 | PTBB-BP0R via BP1C | 2015.4411 | 301.7 | 133.4 |  | 249.73 | 81.43 | 262.91 | 94.61 | PTBB wrt BP0R via BP1C | 2015.4411 | 302.93 | 319.11 | 505.6157
| 57181 | 57187 | BP0U | PTBB | PTB | 57181-57187 | BP0U-PTBB | -574.2 | 0.1 | -583.89 | 0.1 | BP0U-PTBB | 57181-57187 | 87.7 | 74 |  | -574.2 | -560.5 | -583.89 | -570.19 | PTBB-BP0R via BP0U | 2015.4411 | 301.7 | 133.4 |  | 249.64 | 81.34 | 262.79 | 94.49 | PTBB wrt BP0R via BP0U | 2015.4411 | 302.84 | 318.99 | 505.5721
| 57181 | 57187 | BP1C | PTBG | PTB | 57181-57187 | BP1C-PTBG | -502.49 | 0.1 | -519.43 | 0.1 | BP1C-PTBG | 57181-57187 | 280.6 | 46.3 |  | -502.49 | -268.19 | -519.43 | -285.13 | PTBG-BP0R via BP1C | 2015.4411 | 251.4 | 133.4 |  | 196.52 | 78.52 | 216.81 | 98.81 | PTBG wrt BP0R via BP1C | 2015.4411 | 300.02 | 323.31 | 469.1137
| 57181 | 57187 | BP0U | PTBG | PTB | 57181-57187 | BP0U-PTBG | -548.68 | 0.1 | -565.48 | 0.1 | BP0U-PTBG | 57181-57187 | 87.7 | 46.3 |  | -548.68 | -507.28 | -565.48 | -524.08 | PTBG-BP0R via BP0U | 2015.4411 | 251.4 | 133.4 |  | 196.42 | 78.42 | 216.68 | 98.68 | PTBG wrt BP0R via BP0U | 2015.4411 | 299.92 | 323.18 | 469.06
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [6 empty] | The colums below make up Table 3 (Traveling - Reference)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 57001 | 57006 | BP1C | BP0R | BIPM | 57001-57006 | BP1C-BP0R | -59.26 | 0.1 | -55.7 | 0.1 | BP1C-BP0R | 57001-57006 | 257.9 | 270.3 |  | -59.26 | -71.66 | -55.7 | -68.1
| 57127 | 57133 | BP1C | BP0R | BIPM | 57127-57133 | BP1C-BP0R | -59.11 | 0.1 | -55.83 | 0.1 | BP1C-BP0R | 57127-57133 | 261 | 273.7 |  | -59.11 | -71.81 | -55.83 | -68.53
| 57198 | 57205 | BP1C | BP0R | BIPM | 57198-57205 | BP1C-BP0R | -58.94 | 0.2 | -55.73 | 0.1 | BP1C-BP0R | 57198-57205 | 261.1 | 273.7 |  | -58.94 | -71.54 | -55.73 | -68.33
Results for Travel system #1 | [13 empty] | Misclosure |  |  |  | 0.27 |  | 0.43
|  |  | BP1C | BP0R | BIPM | [6 empty] | BP1C-BP0R |  | Mean (chosen value) |  |  |  | -71.67 |  | -68.32
[14 empty] | Mean (arithmetic) |  |  |  | -71.67 |  | -68.32
| 57001 | 57006 | BP0U | BP0R | BIPM | 57001-57006 | BP0U-BP0R | -93.23 | 0.1 | -89.69 | 0.1 | BP0U-BP0R | 57001-57006 | 52.6 | 270.3 |  | -93.23 | -310.93 | -89.69 | -307.39
| 57127 | 57133 | BP0U | BP0R | BIPM | 57127-57133 | BP0U-BP0R | -89.69 | 0.1 | -86.43 | 0.1 | BP0U-BP0R | 57127-57133 | 52.6 | 273.7 |  | -89.69 | -310.79 | -86.43 | -307.53
| 57198 | 57205 | BP0U | BP0R | BIPM | 57198-57205 | BP0U-BP0R | -89.76 | 0.2 | -86.19 | 0.1 | BP0U-BP0R | 57198-57205 | 52.6 | 273.7 |  | -89.76 | -310.86 | -86.19 | -307.29
Results for Travel system #2 | [13 empty] | Misclosure |  |  |  | 0.14 |  | 0.24
|  |  | BP0U | BP0R | BIPM | [6 empty] | BP0U-BP0R |  | Mean (chosen value) |  |  |  | -310.86 |  | -307.4
[14 empty] | Mean (arithmetic) |  |  |  | -310.86 |  | -307.4033

## Phase 3
Phase 3 of the calibration trip is for SU with one traveling systems (BP1K)
Updated 19 October 2015 with results at SU; 10 November 2015 with results of BIPM closure; 17 December 2015 with results based on CGGTTS files
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [5 empty] | Reference delays measured in the actual set-up. | [9 empty] | Measured / known antenna cable delays
[14 empty] | The column "Note" is present to match the Table in the report | [9 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [6 empty] | The colums below make up Table 4 (Traveling - Visited) | [9 empty] | The colums below make up Table 5 (Traveling - Reference) | [8 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
Results based on Rinex data files (10 November 2015) | [30 empty] | BP0R reference values |  | 221.5 | 224.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns) |  |  | Pair | Date | CABDLYV | CABDLYR | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | L1 (ns) | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [6 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | TOTDLY(P3)
Rinex | 57183 | 57188 | BP1K | SU19 | SU | 57183-57188 | BP1K-SU19 | 77.36 | 0.1 | 73.29 | 0.1 | BP1K-SU19 | 57183-57188 | 201.7 | 194.45 |  | 77.36 | 84.61 | 73.29 | 80.54 |  | SU19-BP0R via BP1K | 2015.4452 | 48.17 | 133.4 |  | -336.99 | -251.76 | -337.72 | -252.49 | SU19 wrt BP0R via BP1K | 2015.4452 | -30.26 | -27.99 | -180.0471
Second set-up at SU: Rinex data not affected; CGGTTS data affected by INTDLY values changed to results of absolute calibration. | [27 empty] | -251.525 |  | -252.28
Rinex | 57196 | 57200 | BP1K | SU19 | SU | 57196-57200 | BP1K-SU19 | 77.19 | 0.1 | 73.17 | 0.1 | BP1K-SU19 | 57196-57200 | 201.4 | 194.45 |  | 77.19 | 84.14 | 73.17 | 80.12 |  | SU19-BP0R via BP1K | 2015.4795 | 48.17 | 133.4 |  | -336.52 | -251.29 | -337.3 | -252.07 | SU19 wrt BP0R via BP1K | 2015.4795 | -29.79 | -27.57 | -179.4999
[32 empty] | 2015.4623 | -30.025 | -27.78 | -179.7735
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [6 empty] | The colums below make up Table 3 (Traveling - Reference)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 56897 | 56902 | BP1K | BP0R | BIPM | 56897-56902 | BP1K-BP0R | -27.18 | 0.1 | -31.9 | 0.1 | BP1K-BP0R | 56897-56902 | 45.15 | 270.6 |  | -27.18 | -252.63 | -31.9 | -257.35
| 57332 | 57337 | BP1K | BP0R | BIPM | 57332-57337 | BP1K-BP0R | -14.73 | 0.2 | -19.6 | 0.1 | BP1K-BP0R | 57332-57337 | 51.4 | 288.8 |  | -14.73 | -252.13 | -19.6 | -257
Results for Reference system #1 | [13 empty] | Misclosure |  |  |  | 0.5 |  | 0.35
|  |  | BP1K | BP0R | BIPM | [6 empty] | BP1K-BP0R |  | Mean (chosen value) |  |  |  | -252.38 |  | -257.18
[14 empty] | Mean (arithmetic) |  |  |  | -252.38 |  | -257.175
Results based on CGGTTS data files (17 December 2015). This is indicated for cross-checking results based on Rinex files
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns) |  |  | Pair | Date | CABDLYV | CABDLYR | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | L1 (ns) | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [6 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | TOTDLY(P3)
CGGTTS | 57183 | 57188 | BP1K | SU19 | SU | 57183-57188 | BP1K-SU19 | -9.4 | 0.1 | -13.4 | 0.1 | BP1K-SU19 | 57183-57188 | 0 | 0 | * | -9.4 | -9.4 | -13.4 | -13.4 |  | SU19-BP0R via BP1K | 2015.4452 | 0 | 0 |  | -251.57 | -251.57 | -252.33 | -252.33 | SU19 wrt BP0R via BP1K | 2015.4452 | -30.07 | -27.83
Second set-up at SU: Rinex data not affected; CGGTTS data affected by INTDLY values changed to results of absolute calibration.
CGGTTS | 57196 | 57200 | BP1K | SU19 | SU | 57196-57200 | BP1K-SU19 |  | 0.1 |  | 0.1 | BP1K-SU19 | 57196-57200 | 0 | 0 |  | 0 |  | 0 |  |  | SU19-BP0R via BP1K | 2015.4795 | 0 | 0 | [5 empty] | SU19 wrt BP0R via BP1K | 2015.4795
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [6 empty] | The colums below make up Table 3 (Traveling - Reference)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
CGGTTS | 56897 | 56902 | BP1K | BP0R | BIPM | 56897-56902 | BP1K-BP0R | -261.3 | 0.1 | -265.9 | 0.1 | BP1K-BP0R | 56897-56902 | 0 | 0 |  | -261.3 | -261.3 | -265.9 | -265.9
CGGTTS | 57332 | 57336 | BP1K | BP0R | BIPM | 57332-57336 | BP1K-BP0R | -226.9 | 0.2 | -227.7 | 0.1 | BP1K-BP0R | 57332-57336 | 0 | 0 | * | -226.9 | -260.63 | -227.7 | -265.55 | * INT DLY introduced for BP1K=-33.73/-37.85
Results for Reference system #1 | [13 empty] | Misclosure |  |  |  | 0.67 |  | 0.35
|  |  | BP1K | BP0R | BIPM | [6 empty] | BP1K-BP0R |  | Mean (chosen value) |  |  |  | -260.97 |  | -265.73
[14 empty] | Mean (arithmetic) |  |  |  | -260.965 |  | -265.725

## Phase 2
Phase 2 of the calibration trip has several visits to the reference laboratory (BIPM) and two traveling systemss.
Updated 23 March 2015
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [5 empty] | Reference delays measured in the actual set-up. | [8 empty] | Measured / known antenna cable delays
[14 empty] | The column "Note" is present to match the Table in the report | [8 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [6 empty] | The colums below make up Table 4 (Traveling - Visited) | [8 empty] | The colums below make up Table 5 (Traveling - Reference) | [8 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
[30 empty] | BP0R reference values |  | 221.5 | 224.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | CABDLYV | CABDLYR | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | L1 (ns) | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | TOTDLY(P3)
| 56464 | 56470 | BP1C | PTBB | PTB | 56464-56470 | BP1C-PTBB | -493.38 | 0.1 | -504.14 | 0.1 | BP1C-PTBB | 56464-56470 | 249.4 | 75.3 | * | -493.38 | -319.28 | -504.14 | -330.04 | PTBB-BP0R via BP1C | 2013.4767 | 301.7 | 133.4 |  | 252.39 | 84.09 | 265.3 | 97 | PTBB wrt BP0R via BP1C | 2013.4767 | 305.59 | 321.5 | 507.4091
| 56464 | 56470 | BP0U | PTBB | PTB | 56464-56470 | BP0U-PTBB | -548.75 | 0.1 | -558.47 | 0.1 | BP0U-PTBB | 56464-56470 | 61.5 | 75.3 | * | -548.75 | -562.55 | -558.47 | -572.27 | PTBB-BP0R via BP0U | 2013.4767 | 301.7 | 133.4 |  | 251.65 | 83.35 | 264.69 | 96.39 | PTBB wrt BP0R via BP0U | 2013.4767 | 304.85 | 320.89 | 506.4682
| 56464 | 56470 | BP1C | PTBG | PTB | 56464-56470 | BP1C-PTBG | -467.58 | 0.1 | -485.38 | 0.1 | BP1C-PTBG | 56464-56470 | 249.4 | 48.2 |  | -467.58 | -266.38 | -485.38 | -284.18 | PTBG-BP0R via BP1C | 2013.4767 | 251.4 | 133.4 |  | 199.49 | 81.49 | 219.44 | 101.44 | PTBG wrt BP0R via BP1C | 2013.4767 | 302.99 | 325.94 | 470.7323
| 56464 | 56470 | BP0U | PTBG | PTB | 56464-56470 | BP0U-PTBG | -522.94 | 0.1 | -539.73 | 0.1 | BP0U-PTBG | 56464-56470 | 61.5 | 48.2 |  | -522.94 | -509.64 | -539.73 | -526.43 | PTBG-BP0R via BP0U | 2013.4767 | 251.4 | 133.4 |  | 198.74 | 80.74 | 218.85 | 100.85 | PTBG wrt BP0R via BP0U | 2013.4767 | 302.24 | 325.35 | 469.7351
| 56603 | 56609 | BP1C | TWTF | TL | 56603-56609 | BP1C-TWTF | -319.27 | 0.2 | -323.92 | 0.5 | BP1C-TWTF | 56603-56609 | 234.1 | 52 |  | -319.27 | -137.17 | -323.92 | -141.82 | TWTF-BP0R via BP1C | 2013.8575 | 119.8 | 133.4 |  | 70.28 | 83.88 | 77.08 | 90.68 | TWTF wrt BP0R via BP1C | 2013.8575 | 305.38 | 315.18 | 358.039
| 56603 | 56609 | BP0U | TWTF | TL | 56603-56609 | BP0U-TWTF | -381.23 | 0.2 | -384.57 | 0.4 | BP0U-TWTF | 56603-56609 | 52.6 | 52 |  | -381.23 | -380.63 | -384.57 | -383.97 | TWTF-BP0R via BP0U | 2013.8575 | 119.8 | 133.4 |  | 69.73 | 83.33 | 76.39 | 89.99 | TWTF wrt BP0R via BP0U | 2013.8575 | 304.83 | 314.49 | 357.7053
| 56646 | 56649 | BP1C | TWTF | TL | 56646-56649 | BP1C-TWTF | -319.29 | 0.2 | -323.16 | 0.4 | BP1C-TWTF | 56646-56649 | 234.1 | 52 |  | -319.29 | -137.19 | -323.16 | -141.06 | TWTF-BP0R via BP1C | 2013.9712 | 119.8 | 133.4 |  | 70.3 | 83.9 | 76.32 | 89.92 | TWTF wrt BP0R via BP1C | 2013.9712 | 305.4 | 314.42 | 359.2641
| 56646 | 56649 | BP0U | TWTF | TL | 56646-56649 | BP0U-TWTF | -381.1 | 0.2 | -383.84 | 0.3 | BP0U-TWTF | 56646-56649 | 52.6 | 52 |  | -381.1 | -380.5 | -383.84 | -383.24 | TWTF-BP0R via BP0U | 2013.9712 | 119.8 | 133.4 |  | 69.6 | 83.2 | 75.66 | 89.26 | TWTF wrt BP0R via BP0U | 2013.9712 | 304.7 | 313.76 | 358.5023
| 56715 | 56721 | BP1C | NM0C | NMIJ | 56715-56721 | BP1C-NM0C | -438.12 | 0.1 | -446.04 | 0.1 | BP1C-NM0C | 56715-56721 | 216.7 | 32.4 | * | -438.12 | -253.82 | -446.04 | -261.74 | NM0C-BP0R via BP1C | 2014.1644 | 234.1 | 133.4 |  | 186.93 | 86.23 | 197 | 96.3 | NM0C wrt BP0R via BP1C | 2014.1644 | 307.73 | 320.8 | 489.2369
| 56715 | 56721 | BP0U | NM0C | NMIJ | 56715-56721 | BP0U-NM0C | -497.57 | 0.1 | -504.58 | 0.1 | BP0U-NM0C | 56715-56721 | 33.2 | 32.4 | * | -497.57 | -496.77 | -504.58 | -503.78 | NM0C-BP0R via BP0U | 2014.1644 | 234.1 | 133.4 |  | 185.87 | 85.17 | 196.2 | 95.5 | NM0C wrt BP0R via BP0U | 2014.1644 | 306.67 | 320 | 487.7752
| 56724 | 56731 | BP1C | NC02 | NICT | 56724-56731 | BP1C-NC02 | -416.86 | 0.1 | -418.63 | 0.1 | BP1C-NC02 | 56724-56731 | 666.5 | 429.7 | ** | -416.86 | -180.06 | -418.63 | -181.83 | NC02-BP0R via BP1C | 2014.1904 | 248.5 | 133.4 |  | 113.17 | -1.93 | 117.09 | 1.99 | NC02 wrt BP0R via BP1C | 2014.1904 | 219.57 | 226.49 | 27.6786
| 56724 | 56731 | BP0U | NC02 | NICT | 56724-56731 | BP0U-NC02 | -473.06 | 0.1 | -473.66 | 0.1 | BP0U-NC02 | 56724-56731 | 479.6 | 429.7 | ** | -473.06 | -423.16 | -473.66 | -423.76 | NC02-BP0R via BP0U | 2014.1904 | 248.5 | 133.4 |  | 112.26 | -2.84 | 116.18 | 1.08 | NC02 wrt BP0R via BP0U | 2014.1904 | 218.66 | 225.58 | 26.7686
| 56724 | 56731 | BP1C | SEPA | NICT | 56724-56731 | BP1C-SEPA | -403.12 | 0.1 | -403.58 | 0.2 | BP1C-SEPA | 56724-56731 | 666.5 | 406.1 | ** | -403.12 | -142.72 | -403.58 | -143.18 | SEPA-BP0R via BP1C | 2014.1904 | 213.4 | 133.4 |  | 75.83 | -4.17 | 78.44 | -1.56 | SEPA wrt BP0R via BP1C | 2014.1904 | 217.33 | 222.94 | 15.9625
| 56724 | 56731 | BP0U | SEPA | NICT | 56724-56731 | BP0U-SEPA | -459.29 | 0.1 | -458.62 | 0.2 | BP0U-SEPA | 56724-56731 | 479.6 | 406.1 | ** | -459.29 | -385.79 | -458.62 | -385.12 | SEPA-BP0R via BP0U | 2014.1904 | 213.4 | 133.4 |  | 74.89 | -5.11 | 77.54 | -2.46 | SEPA wrt BP0R via BP0U | 2014.1904 | 216.39 | 222.04 | 14.9608
| 56819 | 56825 | BP1C | IMEJ | NIM | 56819-56825 | BP1C-IMEJ | -78.42 | 0.1 | -76.28 | 0.1 | BP1C-IMEJ | 56819-56825 | 364.6 | 0 | *** | -78.42 | 286.18 | -76.28 | 288.32 | IMEJ-BP0R via BP1C | 2014.4493 | 0 | 133.4 | * | -353.07 | -219.67 | -353.06 | -219.66 | IMEJ wrt BP0R via BP1C | 2014.4493 | 1.83 | 4.84 | -2.8204
| 56819 | 56825 | BP0U | IMEJ | NIM | 56819-56825 | BP0U-IMEJ | -120.94 | 0.1 | -117.47 | 0.1 | BP0U-IMEJ | 56819-56825 | 163.6 | 0 | *** | -120.94 | 42.66 | -117.47 | 46.13 | IMEJ-BP0R via BP0U | 2014.4493 | 0 | 133.4 | * | -353.56 | -220.16 | -353.71 | -220.31 | IMEJ wrt BP0R via BP0U | 2014.4493 | 1.34 | 4.19 | -3.0632
| 56819 | 56825 | BP1C | IMEU | NIM | 56819-56825 | BP1C-IMEU | -185.18 | 0.1 | -193.96 | 0.1 | BP1C-IMEU | 56819-56825 | 364.6 | 115.5 |  | -185.18 | 63.92 | -193.96 | 55.14 | IMEU-BP0R via BP1C | 2014.4493 | 250.3 | 133.4 |  | -130.81 | -247.71 | -119.88 | -236.78 | IMEU wrt BP0R via BP1C | 2014.4493 | -26.21 | -12.28 | 87.0682
| 56819 | 56825 | BP0U | IMEU | NIM | 56819-56825 | BP0U-IMEU | -227.68 | 0.1 | -235.09 | 0.2 | BP0U-IMEU | 56819-56825 | 163.6 | 115.5 |  | -227.68 | -179.58 | -235.09 | -186.99 | IMEU-BP0R via BP0U | 2014.4493 | 250.3 | 133.4 |  | -131.32 | -248.22 | -120.59 | -237.49 | IMEU wrt BP0R via BP0U | 2014.4493 | -26.72 | -12.99 | 86.8671
| 56819 | 56825 | BP1C | BJNM | NIM | 56819-56825 | BP1C-BJNM | 40.25 | 0.2 | 36.79 | 0.1 | BP1C-BJNM | 56819-56825 | 364.6 | 315.3 |  | 40.25 | 89.55 | 36.79 | 86.09 | BJNM-BP0R via BP1C | 2014.4493 | 125 | 133.4 |  | -156.44 | -148.04 | -150.83 | -142.43 | BJNM wrt BP0R via BP1C | 2014.4493 | 73.46 | 82.07 | -130.1424
| 56819 | 56825 | BP0U | BJNM | NIM | 56819-56825 | BP0U-BJNM | -2.25 | 0.1 | -4.41 | 0.1 | BP0U-BJNM | 56819-56825 | 163.6 | 315.3 |  | -2.25 | -153.95 | -4.41 | -156.11 | BJNM-BP0R via BP0U | 2014.4493 | 125 | 133.4 |  | -156.95 | -148.55 | -151.47 | -143.07 | BJNM wrt BP0R via BP0U | 2014.4493 | 72.95 | 81.43 | -130.4516
| 56877 | 56885 | BP1C | PTBB | PTB | 56877-56885 | BP1C-PTBB | -526.42 | 0.1 | -537.28 | 0.1 | BP1C-PTBB | 56877-56885 | 282.5 | 75.3 | * | -526.42 | -319.22 | -537.28 | -330.08 | PTBB-BP0R via BP1C | 2014.611 | 301.7 | 133.4 |  | 252.33 | 84.03 | 265.34 | 97.04 | PTBB wrt BP0R via BP1C | 2014.611 | 305.53 | 321.54 | 507.1945
| 56877 | 56885 | BP0U | PTBB | PTB | 56877-56885 | BP0U-PTBB | -585.51 | 0.1 | -595.15 | 0.1 | BP0U-PTBB | 56877-56885 | 98.2 | 75.3 | * | -585.51 | -562.61 | -595.15 | -572.25 | PTBB-BP0R via BP0U | 2014.611 | 301.7 | 133.4 |  | 251.71 | 83.41 | 264.67 | 96.37 | PTBB wrt BP0R via BP0U | 2014.611 | 304.91 | 320.87 | 506.6518
| 56877 | 56885 | BP1C | PTBG | PTB | 56877-56885 | BP1C-PTBG | -500.71 | 0.1 | -518.69 | 0.1 | BP1C-PTBG | 56877-56885 | 282.5 | 48.2 |  | -500.71 | -266.41 | -518.69 | -284.39 | PTBG-BP0R via BP1C | 2014.611 | 251.4 | 133.4 |  | 199.52 | 81.52 | 219.65 | 101.65 | PTBG wrt BP0R via BP1C | 2014.611 | 303.02 | 326.15 | 470.4841
| 56877 | 56885 | BP0U | PTBG | PTB | 56877-56885 | BP0U-PTBG | -559.75 | 0.1 | -576.56 | 0.1 | BP0U-PTBG | 56877-56885 | 98.2 | 48.2 |  | -559.75 | -509.75 | -576.56 | -526.56 | PTBG-BP0R via BP0U | 2014.611 | 251.4 | 133.4 |  | 198.85 | 80.85 | 218.98 | 100.98 | PTBG wrt BP0R via BP0U | 2014.611 | 302.35 | 325.48 | 469.8142
| 56894 | 56903 | BP1C | RO_4 | ROA | 56894-56903 | BP1C-RO_4 | -218.63 | 0.1 | -218.81 | 0.1 | BP1C-RO_4 | 56894-56903 | 309.3 | 218.9 |  | -218.63 | -128.23 | -218.81 | -128.41 | RO_4-BP0R via BP1C | 2014.6589 | 217.5 | 133.4 |  | 61.34 | -22.76 | 63.67 | -20.43 | RO_4 wrt BP0R via BP1C | 2014.6589 | 198.74 | 204.07 | 189.1051
| 56894 | 56903 | BP0U | RO_4 | ROA | 56894-56903 | BP0U-RO_4 | -278.93 | 0.1 | -277.66 | 0.1 | BP0U-RO_4 | 56894-56903 | 125.7 | 218.9 |  | -278.93 | -372.13 | -277.66 | -370.86 | RO_4-BP0R via BP0U | 2014.6589 | 217.5 | 133.4 |  | 61.23 | -22.87 | 63.28 | -20.82 | RO_4 wrt BP0R via BP0U | 2014.6589 | 198.63 | 203.68 | 189.4278
| 56894 | 56903 | BP1C | RO_5 | ROA | 56894-56903 | BP1C-RO_5 | -22.26 | 0.1 | -15.49 | 0.2 | BP1C-RO_5 | 56894-56903 | 309.3 | 0.3 | *** | -22.26 | 286.74 | -15.49 | 293.51 | RO_5-BP0R via BP1C | 2014.6589 | 0 | 133.4 | * | -353.63 | -220.23 | -358.25 | -224.85 | RO_5 wrt BP0R via BP1C | 2014.6589 | 1.27 | -0.35 | 3.4729
| 56894 | 56903 | BP0U | RO_5 | ROA | 56894-56903 | BP0U-RO_5 | -82.54 | 0.1 | -74.39 | 0.1 | BP0U-RO_5 | 56894-56903 | 125.7 | 0.3 | *** | -82.54 | 42.86 | -74.39 | 51.01 | RO_5-BP0R via BP0U | 2014.6589 | 0 | 133.4 | * | -353.76 | -220.36 | -358.59 | -225.19 | RO_5 wrt BP0R via BP0U | 2014.6589 | 1.14 | -0.69 | 3.6673
| 56894 | 56903 | BP1C | RO_6 | ROA | 56894-56903 | BP1C-RO_6 | 76.37 | 0.1 | 83.41 | 0.1 | BP1C-RO_6 | 56894-56903 | 309.3 | 218.3 |  | 76.37 | 167.37 | 83.41 | 174.41 | RO_6-BP0R via BP1C | 2014.6589 | 66.7 | 133.4 |  | -234.26 | -167.56 | -239.15 | -172.45 | RO_6 wrt BP0R via BP1C | 2014.6589 | 53.94 | 52.05 | -94.7399
| 56894 | 56903 | BP0U | RO_6 | ROA | 56894-56903 | BP0U-RO_6 | 16.14 | 0.1 | 24.54 | 0.1 | BP0U-RO_6 | 56894-56903 | 125.7 | 218.3 |  | 16.14 | -76.46 | 24.54 | -68.06 | RO_6-BP0R via BP0U | 2014.6589 | 66.7 | 133.4 |  | -234.44 | -167.74 | -239.52 | -172.82 | RO_6 wrt BP0R via BP0U | 2014.6589 | 53.76 | 51.68 | -94.6264
| 56894 | 56903 | BP1C | RO_7 | ROA | 56894-56903 | BP1C-RO_7 | 13.6 | 0.1 | 20.04 | 0.1 | BP1C-RO_7 | 56894-56903 | 309.3 | 171.5 |  | 13.6 | 151.4 | 20.04 | 157.84 | RO_7-BP0R via BP1C | 2014.6589 | 81.9 | 133.4 |  | -218.29 | -166.79 | -222.58 | -171.08 | RO_7 wrt BP0R via BP1C | 2014.6589 | 54.71 | 53.42 | -32.897
| 56894 | 56903 | BP0U | RO_7 | ROA | 56894-56903 | BP0U-RO_7 | -46.61 | 0.1 | -38.84 | 0.1 | BP0U-RO_7 | 56894-56903 | 125.7 | 171.5 |  | -46.61 | -92.41 | -38.84 | -84.64 | RO_7-BP0R via BP0U | 2014.6589 | 81.9 | 133.4 |  | -218.49 | -166.99 | -222.94 | -171.44 | RO_7 wrt BP0R via BP0U | 2014.6589 | 54.51 | 53.06 | -32.8497
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [6 empty] | The colums below make up Table 3 (Traveling - Reference)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 56407 | 56412 | BP1C | BP0R | BIPM | 56407-56412 | BP1C-BP0R | 0.5 | 0.1 | 2.35 | 0.1 | BP1C-BP0R | 56407-56412 | 216.6 | 282.6 |  | 0.5 | -65.5 | 2.35 | -63.65
| 56572 | 56579 | BP1C | BP0R | BIPM | 56572-56579 | BP1C-BP0R | -78.44 | 0.1 | -76.26 | 0.1 | BP1C-BP0R | 56572-56579 | 261.7 | 250.9 |  | -78.44 | -67.64 | -76.26 | -65.46
| 56667 | 56673 | BP1C | BP0R | BIPM | 56667-56673 | BP1C-BP0R | -77.68 | 0.2 | -75.53 | 0.1 | BP1C-BP0R | 56667-56673 | 260.9 | 251 |  | -77.68 | -67.78 | -75.53 | -65.63
| 56762 | 56769 | BP1C | BP0R | BIPM | 56762-56769 | BP1C-BP0R | -53.84 | 0.1 | -51.5 | 0.2 | BP1C-BP0R | 56762-56769 | 257.2 | 270.2 |  | -53.84 | -66.84 | -51.5 | -64.5
| 56863 | 56866 | BP1C | BP0R | BIPM | 56863-56866 | BP1C-BP0R | -54.15 | 0.2 | -51.95 | 0.1 | BP1C-BP0R | 56863-56866 | 257.9 | 270.5 |  | -54.15 | -66.75 | -51.95 | -64.55
| 56912 | 56916 | BP1C | BP0R | BIPM | 56912-56916 | BP1C-BP0R | -54.45 | 0.1 | -52.22 | 0.1 | BP1C-BP0R | 56912-56916 | 257.9 | 270.3 |  | -54.45 | -66.85 | -52.22 | -64.62
Results for Travel system #1 | [13 empty] | Misclosure |  |  |  | 2.28 |  | 1.98
|  |  | BP1C | BP0R | BIPM | [6 empty] | BP1C-BP0R |  | Mean (chosen value) |  |  |  | -66.89 |  | -64.74
[14 empty] | Mean (arithmetic) |  |  |  | -66.8933 |  | -64.735
| 56407 | 56412 | BP0U | BP0R | BIPM | 56407-56412 | BP0U-BP0R | -82.39 | 0.1 | -79.3 | 0.2 | BP0U-BP0R | 56407-56412 | 54.2 | 282.6 |  | -82.39 | -310.79 | -79.3 | -307.7
| 56572 | 56579 | BP0U | BP0R | BIPM | 56572-56579 | BP0U-BP0R | -114.31 | 0.2 | -110.9 | 0.2 | BP0U-BP0R | 56572-56579 | 54.2 | 250.9 |  | -114.31 | -311.01 | -110.9 | -307.6
| 56667 | 56673 | BP0U | BP0R | BIPM | 56667-56673 | BP0U-BP0R | -114.46 | 0.1 | -111.11 | 0.1 | BP0U-BP0R | 56667-56673 | 54.2 | 251 |  | -114.46 | -311.26 | -111.11 | -307.91
| 56762 | 56769 | BP0U | BP0R | BIPM | 56762-56769 | BP0U-BP0R | -93.14 | 0.1 | -89.93 | 0.2 | BP0U-BP0R | 56762-56769 | 52.6 | 270.2 |  | -93.14 | -310.74 | -89.93 | -307.53
| 56863 | 56866 | BP0U | BP0R | BIPM | 56863-56866 | BP0U-BP0R | -92.82 | 0.1 | -89.39 | 0.2 | BP0U-BP0R | 56863-56866 | 52.6 | 270.5 |  | -92.82 | -310.72 | -89.39 | -307.29
| 56912 | 56916 | BP0U | BP0R | BIPM | 56912-56916 | BP0U-BP0R | -93.2 | 0.1 | -89.72 | 0.2 | BP0U-BP0R | 56912-56916 | 52.6 | 270.3 |  | -93.2 | -310.9 | -89.72 | -307.42
Results for Travel system #2 | [13 empty] | Misclosure |  |  |  | 0.54 |  | 0.62
|  |  | BP0U | BP0R | BIPM | [6 empty] | BP0U-BP0R |  | Mean (chosen value) |  |  |  | -310.9 |  | -307.58
[14 empty] | Mean (arithmetic) |  |  |  | -310.9033 |  | -307.575

## Phase 1
Phase 1 of the calibration trip is a special case with two references for the closure (BP0R and BP0C), two traveling systems, and two set-ups in the visited laboratory (OPMT/1 and OPMT/2). This sheet should ot be taken as a typical example!
Updated 23 March 2015
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [5 empty] | Reference delays measured in the actual set-up. | [8 empty] | Measured / known antenna cable delays
[14 empty] | The column "Note" is present to match the Table in the report | [8 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2: Raw code differences (Traveling - Local) | [6 empty] | The colums below make up Table 4 (Traveling - Visited) | [8 empty] | The colums below make up Table 5 (Traveling - Reference) | [8 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | L1 (ns) |  | L2 (ns) |  | Pair | Date | CABDLYV | CABDLYR | Note | L1 |  | L2 |  | Pair | Date | L1 (ns) | L2 (ns)
[17 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | TOTDLY(P3)
[30 empty] | BP0R reference values |  | 221.5 | 224.5
First set-up at OP, same reference as OPMT
| 56395 | 56399 | BP0T | OPMT/1 | OP | 56395-56399 | BP0T-OPMT/1 | -354.93 | 0.1 | -364.96 | 0.1 | BP0T-OPMT/1 | 56395-56399 | 50 | 117.2 | * | -354.93 | -422.13 | -364.96 | -432.16 | OPMT/1-BP0R via BP0T | 2013.2849 | 156.5 | 133.4 |  | 108.78 | 85.68 | 118 | 94.9 | OPMT/1 wrt BP0R via BP0T | 2013.2849 | 307.18 | 319.4 | 327.6001
| 56395 | 56399 | BP0U | OPMT/1 | OP | 56395-56399 | BP0U-OPMT/1 | -352.18 | 0.1 | -358.41 | 0.1 | BP0U-OPMT/1 | 56395-56399 | 50 | 117.2 | * | -352.18 | -419.38 | -358.41 | -425.61 | OPMT/1-BP0R via BP0U | 2013.2849 | 156.5 | 133.4 |  | 108.795 | 85.695 | 118.025 | 94.925 | OPMT/1 wrt BP0R via BP0U | 2013.2849 | 307.195 | 319.425 | 327.5996
Second set-up at OP, reference UTC(OP)
| 56400 | 56403 | BP0T | OPMT/2 | OP | 56400-56403 | BP0T-OPMT/2 | -648.03 | 0.1 | -658.04 | 0.1 | BP0T-OPMT/2 | 56400-56403 | 342.6 | 117.2 | * | -648.03 | -422.63 | -658.04 | -432.64 | OPMT/2-BP0R via BP0T | 2013.2973 | 156.5 | 133.4 |  | 109.28 | 86.18 | 118.48 | 95.38 | OPMT/2 wrt BP0R via BP0T | 2013.2973 | 307.68 | 319.88 | 328.131
| 56400 | 56403 | BP0U | OPMT/2 | OP | 56400-56403 | BP0U-OPMT/2 | -645.19 | 0.1 | -651.43 | 0.1 | BP0U-OPMT/2 | 56400-56403 | 342.6 | 117.2 | * | -645.19 | -419.79 | -651.43 | -426.03 | OPMT/2-BP0R via BP0U | 2013.2973 | 156.5 | 133.4 |  | 109.205 | 86.105 | 118.445 | 95.345 | OPMT/2 wrt BP0R via BP0U | 2013.2973 | 307.605 | 319.845 | 327.9942
[30 empty] | BP0C reference values |  | 305.6 | 321.9
[21 empty] | OPMT/1-BP0C via BP0T | 2013.2849 | 156.5 | 234.4 |  | -75.975 | 1.925 | -82.03 | -4.13 | OPMT/1 wrt BP0C via BP0T | 2013.2849 | 307.525 | 317.77
[21 empty] | OPMT/1-BP0C via BP0U | 2013.2849 | 156.5 | 234.4 |  | -75.93 | 1.97 | -81.98 | -4.08 | OPMT/1 wrt BP0C via BP0U | 2013.2849 | 307.57 | 317.82
[21 empty] | OPMT/2-BP0C via BP0T | 2013.2973 | 156.5 | 234.4 |  | -75.475 | 2.425 | -81.55 | -3.65 | OPMT/2 wrt BP0C via BP0T | 2013.2973 | 308.025 | 318.25
[21 empty] | OPMT/2-BP0C via BP0U | 2013.2973 | 156.5 | 234.4 |  | -75.52 | 2.38 | -81.56 | -3.66 | OPMT/2 wrt BP0C via BP0U | 2013.2973 | 307.98 | 318.24
Second part is for the laboratory where the closure takes place (This is a special case with two references for the closure!) | [11 empty] | The colums below make up Table 3 (Traveling - Reference)
| 56380 | 56387 | BP0T | BP0R | BIPM | 56380-56387 | BP0T-BP0R | -84.73 | 0.1 | -85.57 | 0.2 | BP0T-BP0R | 56380-56387 | 54.2 | 282.6 |  | -84.73 | -313.13 | -85.57 | -313.97
| 56407 | 56412 | BP0T | BP0R | BIPM | 56407-56412 | BP0T-BP0R | -85.17 | 0.1 | -85.95 | 0.2 | BP0T-BP0R | 56407-56412 | 54.2 | 282.6 |  | -85.17 | -313.57 | -85.95 | -314.35
[16 empty] | Misclosure |  | 0.44 |  | 0.38
|  |  | BP0T | BP0R | [11 empty] | Mean |  | -313.35 |  | -314.16
| 56380 | 56387 | BP0U | BP0R | BIPM | 56380-56387 | BP0U-BP0R | -81.98 | 0.1 | -79.07 | 0.2 | BP0U-BP0R | 56380-56387 | 54.2 | 282.6 |  | -81.98 | -310.38 | -79.07 | -307.47
| 56407 | 56412 | BP0U | BP0R | BIPM | 56407-56412 | BP0U-BP0R | -82.39 | 0.1 | -79.3 | 0.2 | BP0U-BP0R | 56407-56412 | 54.2 | 282.6 |  | -82.39 | -310.79 | -79.3 | -307.7
[16 empty] | Misclosure |  | 0.41 |  | 0.23
|  |  | BP0U | BP0R | [11 empty] | Mean |  | -310.585 |  | -307.585
| 56380 | 56387 | BP0T | BP0C | BIPM | 56380-56387 | BP0T-BP0C | -463.23 | 0.1 | -479.35 | 0.1 | BP0T-BP0C | 56380-56387 | 54.2 | 89.1 |  | -463.23 | -498.13 | -479.35 | -514.25
| 56407 | 56412 | BP0T | BP0C | BIPM | 56407-56412 | BP0T-BP0C | -463.38 | 0.1 | -479.43 | 0.1 | BP0T-BP0C | 56407-56412 | 54.2 | 88.9 |  | -463.38 | -498.08 | -479.43 | -514.13
[16 empty] | Misclosure |  | 0.05 |  | 0.12
|  |  | BP0T | BP0C | [11 empty] | Mean |  | -498.105 |  | -514.19
| 56380 | 56387 | BP0U | BP0C | BIPM | 56380-56387 | BP0U-BP0C | -460.46 | 0.1 | -472.84 | 0.1 | BP0U-BP0C | 56380-56387 | 54.2 | 89.1 |  | -460.46 | -495.36 | -472.84 | -507.74
| 56407 | 56412 | BP0U | BP0C | BIPM | 56407-56412 | BP0U-BP0C | -460.56 | 0.1 | -472.74 | 0.1 | BP0U-BP0C | 56407-56412 | 54.2 | 88.9 |  | -460.56 | -495.26 | -472.74 | -507.44
[16 empty] | Misclosure |  | 0.1 |  | 0.3
|  |  | BP0U | BP0C | [11 empty] | Mean |  | -495.31 |  | -507.59