## Phase 3
Phase 3 of the calibration trip is BIPM-NIST-USNO(possibly)-BIPM and two traveling systems.
Updated 27/08/2021 for opening at BIPM; 19/10, 29/10 for NIST preliminary; XX/XX/2022 for USNO preliminary; 04/05/2022 final.
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [25 empty] | Reference delays measured in the actual set-up. | [28 empty] | Measured / known antenna cable delays
[34 empty] | The column "Note" is present to match the Table in the report | [28 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [26 empty] | The colums below make up Table 4 (Traveling - Visited) | [28 empty] | The colums below make up Table 5 (Visited - Reference) | [28 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known) | [11 empty] | BDS3 values from BP27 absolute calibration and 2021-08 BIPM comparison
[90 empty] | BP21 reference values 1001-2020 |  | 28.4 | 27.3 | 30.6 | 0 | 30.7 | 30.9 | 27.7 | 28.6 | 24.4 | 25.5 | 30.4 | 30.3
| Start | End | Travel | Local | Labo | Date | Pair | RDIF(P1) | Unc | RDIF(P2) | Unc | RDIF(C1) | Unc | RDIF(C2) | Unc | RDIF(E1) | Unc | RDIF(E5a) | Unc | RDIF(D1) | Unc | RDIF(D2) | Unc | RDIF(B1) | Unc | RDIF(B2I) | Unc | RDIF(B1C) | Unc | RDIF(B2a) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a |  | D1 |  | D2 |  | B1 |  | B2 |  | BC |  | B5 |  | Pair | Date | CABDLYV | CABDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | D2 |  | B1 |  | B2 |  | BC |  | B5 |  | Pair | Date | P1 (ns) | P2 (ns) | C1 (ns) | C2 (ns) | E1 (ns) | E5 (ns) | D1(ns) | D2(ns) | B1(ns) | B2(ns) | BC(ns) | B5(ns)
[37 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY |  |  | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | INTDLYV | [9 empty] | Type | (CAB-REF)apriori | TOTDLY(P3)
after step | 59491 | 59498 | BP1J | NISG | NIST | 59491-59498 | BP1J-NISG | 866.86 | 0.1 | 868.35 | 0.1 | 865.88 | 0.1 | 867.52 | 0.1 | 865.39 | 0.1 | 875.69 | 0.1 | 880.43 | 0.5 | 877.17 | 0.2 | [8 empty] | BP1J-NISG | 59491-59498 | 605.4 | 1592.2 |  | 866.86 | -119.94 | 868.35 | -118.45 | 865.88 | -120.92 | 867.52 | -119.28 | 865.39 | -121.41 | 875.69 | -111.11 | 880.43 | -106.37 | 877.17 | -109.63 | [8 empty] | NISG-BP21 via BP1J | 2021.7712 | 298.5 | 140.8 |  | 180.455 | 22.755 | 179.855 | 22.155 | 180.64 | 22.94 |  |  | 180.535 | 22.835 | 180.055 | 22.355 | 179.52 | 21.82 | 180.745 | 23.045 | [8 empty] | NISG wrt BP21 via BP1J | 2021.7712 | 51.155 | 49.455 | 53.54 |  | 53.535 | 53.255 | 49.52 | 51.645 |  |  |  |  | PolaRx5 |  | -1239.9185
| 59491 | 59498 | BP25 | NISG | NIST | 59491-59498 | BP25-NISG | 988.09 | 0.1 | 994 | 0.1 | 986.76 | 0.1 | 992.88 | 0.1 | 986.27 | 0.1 | 992.59 | 0.1 | 991.17 | 0.2 | 995.26 | 0.1 | 995.09 | 0.2 | 995.84 | 0.3 | 985.99 | 0.2 | 992.46 | 0.2 | BP25-NISG | 59491-59498 | 466.5 | 1592.2 |  | 988.09 | -137.61 | 994 | -131.7 | 986.76 | -138.94 | 992.88 | -132.82 | 986.27 | -139.43 | 992.59 | -133.11 | 991.17 | -134.53 | 995.26 | -130.44 | 995.09 | -130.61 | 995.84 | -129.86 |  |  |  |  | NISG-BP21 via BP25 | 2021.7712 | 298.5 | 140.8 |  | 159.795 | 2.095 | 159.445 | 1.745 | 159.77 | 2.07 |  |  | 159.865 | 2.165 | 159.59 | 1.89 | 158.775 | 1.075 | 160.135 | 2.435 | 159.525 | 1.825 | 158.955 | 1.255 |  |  |  |  | NISG wrt BP21 via BP25 | 2021.7712 | 30.495 | 29.045 | 32.67 |  | 32.865 | 32.79 | 28.775 | 31.035 | 26.225 | 26.755 |  |  | PolaRx5 |  | -1260.9648
before step | 59491 | 59492 | BP1J | NISG | NIST | 59491-59492 | BP1J-NISG | 886.92 | 0.1 | 888.41 | 0.1 | 885.93 | 0.1 | 887.58 | 0.1 | 885.47 | 0.1 | 895.77 | 0.1 | 900.51 | 0.5 | 897.23 | 0.2 | [8 empty] | BP1J-NISG | 59491-59492 | 605.4 | 1592.2 |  | 886.92 | -99.88 | 888.41 | -98.39 | 885.93 | -100.87 | 887.58 | -99.22 | 885.47 | -101.33 | 895.77 | -91.03 | 900.51 | -86.29 | 897.23 | -89.57 | [8 empty] | NISG-BP21 via BP1J | 2021.763 | 298.5 | 140.8 |  | 160.395 | 2.695 | 159.795 | 2.095 | 160.59 | 2.89 |  |  | 160.455 | 2.755 | 159.975 | 2.275 | 159.44 | 1.74 | 160.685 | 2.985 | [8 empty] | NISG wrt BP21 via BP1J | 2021.763 | 31.095 | 29.395 | 33.49 |  | 33.455 | 33.175 | 29.44 | 31.585 |  |  |  |  | PolaRx6
after step | 59494 | 59498 | BP1J | NIST | NIST | 59494-59498 | BP1J-NIST | -532.68 | 0.1 | -533.82 | 0.1 | -532.11 | 0.1 | [18 empty] | BP1J-NIST | 59494-59498 | 605.4 | 65.9 |  | -532.68 | 6.82 | -533.82 | 5.68 | -532.11 | 7.39 | [18 empty] | NIST-BP21 via BP1J | 2021.7753 | 275.5 | 140.8 |  | 53.695 | -81.005 | 55.725 | -78.975 | 52.33 | -82.37 | [18 empty] | NIST wrt BP21 via BP1J | 2021.7753 | -52.605 | -51.675 | -51.77 | [9 empty] | Novatel |  | 155.5581
| 59494 | 59498 | BP25 | NIST | NIST | 59494-59498 | BP25-NIST | -411.45 | 0.1 | -408.26 | 0.1 | -411.21 | 0.1 | [18 empty] | BP25-NIST | 59494-59498 | 466.5 | 65.9 |  | -411.45 | -10.85 | -408.26 | -7.66 | -411.21 | -10.61 | [18 empty] | NIST-BP21 via BP25 | 2021.7753 | 275.5 | 140.8 |  | 33.035 | -101.665 | 35.405 | -99.295 | 31.44 | -103.26 | [18 empty] | NIST wrt BP21 via BP25 | 2021.7753 | -73.265 | -71.995 | -72.66 | [9 empty] | Novatel |  | 134.3728
before step | 59491 | 59492 | BP1J | NIST | NIST | 59491-59492 | BP1J-NIST | -512.72 | 0.1 | -513.94 | 0.1 | -512.15 | 0.1 | [18 empty] | BP1J-NIST | 59491-59492 | 605.4 | 65.9 |  | -512.72 | 26.78 | -513.94 | 25.56 | -512.15 | 27.35 | [18 empty] | NIST-BP21 via BP1J | 2021.763 | 275.5 | 140.8 |  | 33.735 | -100.965 | 35.845 | -98.855 | 32.37 | -102.33 | [18 empty] | NIST wrt BP21 via BP1J | 2021.763 | -72.565 | -71.555 | -71.73 | [9 empty] | Novatel |  | 135.4746
[6 empty] | 00000-00000 | 0000-0000 | [24 empty] | 0000-0000 | 00000-00000 | [72 empty] | 0
after step | 59510 | 59514 | BP1J | NISG | NIST | 59510-59514 | BP1J-NISG | 887.59 | 0.1 | 889.17 | 0.1 | 886.63 | 0.1 | 888.32 | 0.1 | 886.16 | 0.1 | 896.53 | 0.1 | 901.21 | 0.5 | 897.87 | 0.2 | [8 empty] | BP1J-NISG | 59510-59514 | 606.1 | 1592.2 |  | 887.59 | -98.51 | 889.17 | -96.93 | 886.63 | -99.47 | 888.32 | -97.78 | 886.16 | -99.94 | 896.53 | -89.57 | 901.21 | -84.89 | 897.87 | -88.23 | [8 empty] | NISG-BP21 via BP1J | 2021.8192 | 298.5 | 140.8 |  | 159.025 | 1.325 | 158.335 | 0.635 | 159.19 | 1.49 |  |  | 159.065 | 1.365 | 158.515 | 0.815 | 158.04 | 0.34 | 159.345 | 1.645 | [8 empty] | NISG wrt BP21 via BP1J | 2021.8192 | 29.725 | 27.935 | 32.09 |  | 32.065 | 31.715 | 28.04 | 30.245 |  |  |  |  | PolaRx5 |  | -1261.2095
| 59510 | 59514 | BP25 | NISG | NIST | 59510-59514 | BP25-NISG | 988.64 | 0.1 | 994.52 | 0.1 | 987.27 | 0.1 | 993.39 | 0.1 | 986.82 | 0.1 | 993.26 | 0.1 | 991.64 | 0.2 | 995.8 | 0.1 | 995.64 | 0.2 | 996.42 | 0.3 | 986.52 | 0.2 | 993.13 | 0.2 | BP25-NISG | 59510-59514 | 467.17 | 1592.2 |  | 988.64 | -136.39 | 994.52 | -130.51 | 987.27 | -137.76 | 993.39 | -131.64 | 986.82 | -138.21 | 993.26 | -131.77 | 991.64 | -133.39 | 995.8 | -129.23 | 995.64 | -129.39 | 996.42 | -128.61 |  |  |  |  | NISG-BP21 via BP25 | 2021.8192 | 298.5 | 140.8 |  | 158.575 | 0.875 | 158.255 | 0.555 | 158.59 | 0.89 |  |  | 158.645 | 0.945 | 158.25 | 0.55 | 157.635 | -0.065 | 158.925 | 1.225 | 158.305 | 0.605 | 157.705 | 0.005 |  |  |  |  | NISG wrt BP21 via BP25 | 2021.8192 | 29.275 | 27.855 | 31.49 |  | 31.645 | 31.45 | 27.635 | 29.825 | 25.005 | 25.505 |  |  | PolaRx5 |  | -1262.2311
| 59510 | 59514 | BP1J | NIST | NIST | 59510-59514 | BP1J-NIST | -512.47 | 0.1 | -513.37 | 0.1 | -511.87 | 0.1 | [18 empty] | BP1J-NIST | 59510-59514 | 606.1 | 65.9 |  | -512.47 | 27.73 | -513.37 | 26.83 | -511.87 | 28.33 | [18 empty] | NIST-BP21 via BP1J | 2021.8192 | 275.5 | 140.8 |  | 32.785 | -101.915 | 34.575 | -100.125 | 31.39 | -103.31 | [18 empty] | NIST wrt BP21 via BP1J | 2021.8192 | -73.515 | -72.825 | -72.71 | [10 empty] | 0 | 135.019
| 59510 | 59514 | BP25 | NIST | NIST | 59510-59514 | BP25-NIST | -411.44 | 0.1 | -408.05 | 0.1 | -411.18 | 0.1 | [18 empty] | BP25-NIST | 59510-59514 | 467.17 | 65.9 |  | -411.44 | -10.17 | -408.05 | -6.78 | -411.18 | -9.91 | [18 empty] | NIST-BP21 via BP25 | 2021.8192 | 275.5 | 140.8 |  | 32.355 | -102.345 | 34.525 | -100.175 | 30.74 | -103.96 | [18 empty] | NIST wrt BP21 via BP25 | 2021.8192 | -73.945 | -72.875 | -73.36 | [10 empty] | 0 | 134.0018
| 59510 | 59514 | BP1J | NISS | NIST | 59510-59514 | BP1J-NISS | -421.52 | 0.2 | -422.21 | 0.1 | -421.96 | 0.2 | -422.12 | 0.1 |  |  |  |  | -420.07 | 0.1 | -413.21 | 0.1 | [8 empty] | BP1J-NISS | 59510-59514 | 606.1 | 299.4 |  | -421.52 | -114.82 | -422.21 | -115.51 | -421.96 | -115.26 | -422.12 | -115.42 |  |  |  |  | -420.07 | -113.37 | -413.21 | -106.51 | [8 empty] | NISS-BP21 via BP1J | 2021.8192 | 298.9 | 140.8 |  | 175.335 | 17.235 | 176.915 | 18.815 | 174.98 | 16.88 | [6 empty] | 186.52 | 28.42 | 177.625 | 19.525 | [8 empty] | NISS wrt BP21 via BP1J | 2021.8192 | 45.635 | 46.115 | 47.48 |  |  |  | 56.12 | 48.125 | [6 empty] | 44.3934
| 59510 | 59514 | BP25 | NISS | NIST | 59510-59514 | BP25-NISS | -320.46 | 0.2 | -316.82 | 0.1 | -321.29 | 0.2 | -317.03 | 0.1 |  |  |  |  | -329.26 | 0.5 | -315.39 | 0.3 | [8 empty] | BP25-NISS | 59510-59514 | 467.17 | 299.4 |  | -320.46 | -152.69 | -316.82 | -149.05 | -321.29 | -153.52 | -317.03 | -149.26 |  |  |  |  | -329.26 | -161.49 | -315.39 | -147.62 | [8 empty] | NISS-BP21 via BP25 | 2021.8192 | 298.9 | 140.8 |  | 174.875 | 16.775 | 176.795 | 18.695 | 174.35 | 16.25 | [6 empty] | 185.735 | 27.635 | 177.315 | 19.215 | [8 empty] | NISS wrt BP21 via BP25 | 2021.8192 | 45.175 | 45.995 | 46.85 |  |  |  | 55.335 | 47.815 | [6 empty] | 43.4081
| 59510 | 59514 | BP1J | NISX | NIST | 59510-59514 | BP1J-NISX | -137.45 |  | -135.58 |  | -138.33 |  | -136.55 |  | -138.91 |  | -129.45 | [13 empty] | BP1J-NISX | 59510-59514 | 606.1 | 452.29 |  | -137.45 | 16.36 | -135.58 | 18.23 | -138.33 | 15.48 | -136.55 | 17.26 | -138.91 | 14.9 | -129.45 | 24.36 | [12 empty] | NISX-BP21 via BP1J | 2021.8192 | 185 | 140.8 |  | 44.155 | -0.045 | 43.175 | -1.025 | 44.24 | 0.04 |  |  | 44.225 | 0.025 | 44.585 | 0.385 | [12 empty] | NISX wrt BP21 via BP1J | 2021.8192 | 28.355 | 26.275 | 30.64 |  | 30.725 | 31.285 | [6 empty] | PolaRx5 |  | -235.7214
| 59510 | 59514 | BP25 | NISX | NIST | 59510-59514 | BP25-NISX | -36.44 |  | -30.22 |  | -37.69 |  | -31.48 |  | -38.28 |  | -32.75 | [13 empty] | BP25-NISX | 59510-59514 | 467.17 | 452.29 |  | -36.44 | -21.56 | -30.22 | -15.34 | -37.69 | -22.81 | -31.48 | -16.6 | -38.28 | -23.4 | -32.75 | -17.87 | [12 empty] | NISX-BP21 via BP25 | 2021.8192 | 185 | 140.8 |  | 43.745 | -0.455 | 43.085 | -1.115 | 43.64 | -0.56 |  |  | 43.835 | -0.365 | 44.35 | 0.15 | [12 empty] | NISX wrt BP21 via BP25 | 2021.8192 | 27.945 | 26.185 | 30.04 |  | 30.335 | 31.05 | [6 empty] | PolaRx5 |  | -236.6258
Used for final report with 20 days | 59510 | 59529 | BP1J | NISG | NIST | 59510-59529 | BP1J-NISG | 887.62 | 0.1 | 889.17 | 0.1 | 886.65 | 0.1 | 888.33 | 0.1 | 886.18 | 0.1 | 896.53 | 0.1 | 901.27 | 0.5 | 897.85 | 0.2 | [8 empty] | BP1J-NISG | 59510-59529 | 606.1 | 1592.2 |  | 887.62 | -98.48 | 889.17 | -96.93 | 886.65 | -99.45 | 888.33 | -97.77 | 886.18 | -99.92 | 896.53 | -89.57 | 901.27 | -84.83 | 897.85 | -88.25 | [8 empty] | NISG-BP21 via BP1J | 2021.8397 | 298.5 | 140.8 |  | 158.995 | 1.295 | 158.335 | 0.635 | 159.17 | 1.47 |  |  | 159.045 | 1.345 | 158.515 | 0.815 | 157.98 | 0.28 | 159.365 | 1.665 | [8 empty] | NISG wrt BP21 via BP1J | 2021.8397 | 29.695 | 27.935 | 32.07 |  | 32.045 | 31.715 | 27.98 | 30.265 |  |  |  |  | PolaRx5 |  | -1261.2858
| 59510 | 59529 | BP25 | NISG | NIST | 59510-59529 | BP25-NISG | 988.65 | 0.1 | 994.52 | 0.1 | 987.31 | 0.1 | 993.43 | 0.1 | 986.83 | 0.1 | 993.37 | 0.1 | 991.69 | 0.2 | 995.75 | 0.1 | 995.65 | 0.2 | 996.5 | 0.3 | 986.53 | 0.2 | 993.23 | 0.2 | BP25-NISG | 59510-59529 | 467.19 | 1592.2 |  | 988.65 | -136.36 | 994.52 | -130.49 | 987.31 | -137.7 | 993.43 | -131.58 | 986.83 | -138.18 | 993.37 | -131.64 | 991.69 | -133.32 | 995.75 | -129.26 | 995.65 | -129.36 | 996.5 | -128.51 | 986.53 | -138.48 | 993.23 | -131.78 | NISG-BP21 via BP25 | 2021.8397 | 298.5 | 140.8 |  | 158.545 | 0.845 | 158.235 | 0.535 | 158.53 | 0.83 |  |  | 158.615 | 0.915 | 158.12 | 0.42 | 157.565 | -0.135 | 158.955 | 1.255 | 158.275 | 0.575 | 157.605 | -0.095 | 158.6 | 0.9 | 157.88 | 0.18 | NISG wrt BP21 via BP25 | 2021.8397 | 29.245 | 27.835 | 31.43 |  | 31.615 | 31.32 | 27.565 | 29.855 | 24.975 | 25.405 | 31.3 | 30.48 | PolaRx5 |  | -1262.2765
| 59510 | 59529 | BP1J | NIST | NIST | 59510-59529 | BP1J-NIST | -512.46 | 0.1 | -513.32 | 0.1 | -511.84 | 0.1 | [18 empty] | BP1J-NIST | 59510-59529 | 606.1 | 65.9 |  | -512.46 | 27.74 | -513.32 | 26.88 | -511.84 | 28.36 | [18 empty] | NIST-BP21 via BP1J | 2021.8397 | 275.5 | 140.8 |  | 32.775 | -101.925 | 34.525 | -100.175 | 31.36 | -103.34 | [18 empty] | NIST wrt BP21 via BP1J | 2021.8397 | -73.525 | -72.875 | -72.74 | [9 empty] | Novatel | 0 | 135.0708
| 59510 | 59529 | BP25 | NIST | NIST | 59510-59529 | BP25-NIST | -411.42 | 0.1 | -407.99 | 0.1 | -411.15 | 0.1 | [18 empty] | BP25-NIST | 59510-59529 | 467.19 | 65.9 |  | -411.42 | -10.13 | -407.99 | -6.7 | -411.15 | -9.86 | [18 empty] | NIST-BP21 via BP25 | 2021.8397 | 275.5 | 140.8 |  | 32.315 | -102.385 | 34.445 | -100.255 | 30.69 | -104.01 | [18 empty] | NIST wrt BP21 via BP25 | 2021.8397 | -73.985 | -72.955 | -73.41 | [9 empty] | Novatel | 0 | 134.0237
| 59510 | 59529 | BP1J | NISS | NIST | 59510-59529 | BP1J-NISS | -421.3 | 0.2 | -422.04 | 0.1 | -421.73 | 0.2 | -421.92 | 0.1 |  |  |  |  | -419.92 | 0.1 | -413.01 | 0.1 | [8 empty] | BP1J-NISS | 59510-59529 | 606.1 | 299.4 |  | -421.3 | -114.6 | -422.04 | -115.34 | -421.73 | -115.03 | -421.92 | -115.22 |  |  |  |  | -419.92 | -113.22 | -413.01 | -106.31 | [8 empty] | NISS-BP21 via BP1J | 2021.8397 | 298.9 | 140.8 |  | 175.115 | 17.015 | 176.745 | 18.645 | 174.75 | 16.65 | [6 empty] | 186.37 | 28.27 | 177.425 | 19.325 | [8 empty] | NISS wrt BP21 via BP1J | 2021.8397 | 45.415 | 45.945 | 47.25 |  |  |  | 55.97 | 47.925 |  |  |  |  | PolaRx3 |  | 44.0961
| 59510 | 59529 | BP25 | NISS | NIST | 59510-59529 | BP25-NISS | -320.26 | 0.2 | -316.66 | 0.1 | -321.06 | 0.2 | -316.82 | 0.1 |  |  |  |  | -329.18 | 0.5 | -315.28 | 0.3 | [8 empty] | BP25-NISS | 59510-59529 | 467.19 | 299.4 |  | -320.26 | -152.47 | -316.66 | -148.87 | -321.06 | -153.27 | -316.82 | -149.03 |  |  |  |  | -329.18 | -161.39 | -315.28 | -147.49 | [8 empty] | NISS-BP21 via BP25 | 2021.8397 | 298.9 | 140.8 |  | 174.655 | 16.555 | 176.615 | 18.515 | 174.1 | 16 | [6 empty] | 185.635 | 27.535 | 177.185 | 19.085 | [8 empty] | NISS wrt BP21 via BP25 | 2021.8397 | 44.955 | 45.815 | 46.6 |  |  |  | 55.235 | 47.685 |  |  |  |  | PolaRx3 |  | 43.1263
| 59510 | 59529 | BP1J | NISX | NIST | 59510-59529 | BP1J-NISX | -137.45 | 0.1 | -135.69 | 0.1 | -138.33 | 0.1 | -136.58 | 0.1 | -138.91 | 0.1 | -129.49 | 0.1 | -126.57 | 0.5 | -128.27 | 0.3 | [8 empty] | BP1J-NISX | 59510-59529 | 606.1 | 452.3 |  | -137.45 | 16.35 | -135.69 | 18.11 | -138.33 | 15.47 | -136.58 | 17.22 | -138.91 | 14.89 | -129.49 | 24.31 | [12 empty] | NISX-BP21 via BP1J | 2021.8397 | 185 | 140.8 |  | 44.165 | -0.035 | 43.295 | -0.905 | 44.25 | 0.05 |  |  | 44.235 | 0.035 | 44.635 | 0.435 | [12 empty] | NISX wrt BP21 via BP1J | 2021.8397 | 28.365 | 26.395 | 30.65 |  | 30.735 | 31.335 | [6 empty] | PolaRx5 |  | -235.8914
| 59510 | 59529 | BP25 | NISX | NIST | 59510-59529 | BP25-NISX | -36.45 | 0.1 | -30.26 | 0.1 | -37.68 | 0.1 | -31.48 | 0.1 | -38.29 | 0.1 | -32.68 | 0.1 | -36.3 | 0.3 | -30.47 | 0.1 | -30.27 | 0.1 | -30.07 | 0.3 | -38.72 | 0.2 | -33.04 | 0.1 | BP25-NISX | 59510-59529 | 467.19 | 452.3 |  | -36.45 | -21.56 | -30.26 | -15.37 | -37.68 | -22.79 | -31.48 | -16.59 | -38.29 | -23.4 | -32.68 | -17.79 | [12 empty] | NISX-BP21 via BP25 | 2021.8397 | 185 | 140.8 |  | 43.745 | -0.455 | 43.115 | -1.085 | 43.62 | -0.58 |  |  | 43.835 | -0.365 | 44.27 | 0.07 | [12 empty] | NISX wrt BP21 via BP25 | 2021.8397 | 27.945 | 26.215 | 30.02 |  | 30.335 | 30.97 | [6 empty] | PolaRx5 |  | -236.6822
Used for final report: End 2021 | 59569 | 59579 | BP1J | USN6 | USNO | 59569-59579 | BP1J-USN6 | 64.87 | 0.1 | 67.16 | 0.1 | 65.42 | 0.1 | [6 empty] | 25.63 | 0.1 | 21.24 | 0.5 | [8 empty] | BP1J-USN6 | 59569-59579 | 171.38 | 0 |  | 64.87 | 236.25 | 67.16 | 238.54 | 65.42 | 236.8 | [6 empty] | 25.63 | 197.01 | 21.24 | 192.62 | [8 empty] | USN6-BP21 via BP1J | 2021.989 | 0 | 140.8 |  | -175.735 | -34.935 | -177.135 | -36.335 | -177.08 | -36.28 | [6 empty] | -123.86 | 16.94 | -121.505 | 19.295 | [8 empty] | USN6 wrt BP21 via BP1J | 2021.989 | -6.535 | -9.035 | -5.68 |  |  |  | 44.64 | 47.895 |  |  |  |  | Novatel |  | -2.6725
| 59569 | 59579 | BP25 | USN6 | USNO | 59569-59579 | BP25-USN6 | [24 empty] | BP25-USN6 | 59569-59579 |  | 0 | [25 empty] | USN6-BP21 via BP25 | 2021.989 | 0 | 140.8 | [25 empty] | USN6 wrt BP21 via BP25 | 2021.989 | [14 empty] | 0
| 59569 | 59579 | BP1J | USN7 | USNO | 59569-59579 | BP1J-USN7 | -94.95 | 0.1 | -91.19 | 0.1 | -95.82 | 0.1 | -92.31 | 0.1 | -96.39 | 0.1 | -87.85 | 0.1 | -82.75 | 0.5 | -83.93 | 0.2 | [8 empty] | BP1J-USN7 | 59569-59579 | 171.38 | 0 |  | -94.95 | 76.43 | -91.19 | 80.19 | -95.82 | 75.56 | -92.31 | 79.07 | -96.39 | 74.99 | -87.85 | 83.53 | -82.75 | 88.63 | -83.93 | 87.45 | [8 empty] | USN7-BP21 via BP1J | 2021.989 | 0 | 140.8 |  | -15.915 | 124.885 | -18.785 | 122.015 | -15.84 | 124.96 |  |  | -15.865 | 124.935 | -14.585 | 126.215 | -15.48 | 125.32 | -16.335 | 124.465 | [8 empty] | USN7 wrt BP21 via BP1J | 2021.989 | 153.285 | 149.315 | 155.56 |  | 155.635 | 157.115 | 153.02 | 153.065 |  |  |  |  | PolaRx5 |  | 159.4187
| 59569 | 59579 | BP25 | USN7 | USNO | 59569-59579 | BP25-USN7 | [24 empty] | BP25-USN7 | 59569-59579 |  | 0 | [25 empty] | USN7-BP21 via BP25 | 2021.989 | 0 | 140.8 | [25 empty] | USN7 wrt BP21 via BP25 | 2021.989 | [14 empty] | 0
| 59569 | 59579 | BP1J | USN8 | USNO | 59569-59579 | BP1J-USN8 | -141.95 | 0.1 | -139.24 | 0.1 | -142.86 | 0.1 | -140.33 | 0.1 | -143.45 | 0.1 | -135.12 | 0.1 | -131.75 | 0.5 | -131.3 | 0.2 | [8 empty] | BP1J-USN8 | 59569-59579 | 171.38 | 0 |  | -141.95 | 29.43 | -139.24 | 32.14 | -142.86 | 28.52 | -140.33 | 31.05 | -143.45 | 27.93 | -135.12 | 36.26 | -131.75 | 39.63 | -131.3 | 40.08 | [8 empty] | USN8-BP21 via BP1J | 2021.989 | 0 | 140.8 |  | 31.085 | 171.885 | 29.265 | 170.065 | 31.2 | 172 |  |  | 31.195 | 171.995 | 32.685 | 173.485 | 33.52 | 174.32 | 31.035 | 171.835 | [8 empty] | USN8 wrt BP21 via BP1J | 2021.989 | 200.285 | 197.365 | 202.6 |  | 202.695 | 204.385 | 202.02 | 200.435 |  |  |  |  | PolaRx5 |  | 204.7964
| 59569 | 59579 | BP25 | USN8 | USNO | 59569-59579 | BP25-USN8 | [24 empty] | BP25-USN8 | 59569-59579 |  | 0 | [25 empty] | USN8-BP21 via BP25 | 2021.989 | 0 | 140.8 | [25 empty] | USN8 wrt BP21 via BP25 | 2021.989 | [14 empty] | 0
| 59569 | 59579 | BP1J | USN9 | USNO | 59569-59579 | BP1J-USN9 | -143.95 | 0.1 | -161.51 | 0.1 | -143.36 | 0.1 | -159.46 | 0.1 | -144.92 | 0.1 | -159.77 | 0.1 | -161.83 | 0.2 | -153.72 | 0.2 | [8 empty] | BP1J-USN9 | 59569-59579 | 171.38 | 0 |  | -143.95 | 27.43 | -161.51 | 9.87 | -143.36 | 28.02 | -159.46 | 11.92 | -144.92 | 26.46 | -159.77 | 11.61 | -161.83 | 9.55 | -153.72 | 17.66 | [8 empty] | USN9-BP21 via BP1J | 2021.989 | 0 | 140.8 |  | 33.085 | 173.885 | 51.535 | 192.335 | 31.7 | 172.5 |  |  | 32.665 | 173.465 | 57.335 | 198.135 | 63.6 | 204.4 | 53.455 | 194.255 | [8 empty] | USN9 wrt BP21 via BP1J | 2021.989 | 202.285 | 219.635 | 203.1 |  | 204.165 | 229.035 | 232.1 | 222.855 |  |  |  |  | Novatel |  | 175.4793
| 59569 | 59579 | BP25 | USN9 | USNO | 59569-59579 | BP25-USN9 | [24 empty] | BP25-USN9 | 59569-59579 |  | 0 | [25 empty] | USN9-BP21 via BP25 | 2021.989 | 0 | 140.8 | [25 empty] | USN9 wrt BP21 via BP25 | 2021.989 | [14 empty] | 0
Begin 2022 (not used in final report, but consistent) | 59580 | 59586 | BP1J | USN6 | USNO | 59580-59586 | BP1J-USN6 | 64.79 | 0.1 | 67.42 | 0.1 | 65.37 | 0.1 | [6 empty] | 25.54 | 0.1 | 21.18 | 0.5 | [8 empty] | BP1J-USN6 | 59580-59586 | 171.4 | 0 |  | 64.79 | 236.19 | 67.42 | 238.82 | 65.37 | 236.77 | [6 empty] | 25.54 | 196.94 | 21.18 | 192.58 | [8 empty] | USN6-BP21 via BP1J | 2022.0137 |  | 140.8 |  | -175.675 | -34.875 | -177.415 | -36.615 | -177.05 | -36.25 | [6 empty] | -123.79 | 17.01 | -121.465 | 19.335 | [8 empty] | USN6 wrt BP21 via BP1J | 2022.0137 | -6.475 | -9.315 | -5.65 |  |  |  | 44.71 | 47.935 |  |  |  |  | Novatel |  | -2.0872
| 59580 | 59586 | BP25 | USN6 | USNO | 59580-59586 | BP25-USN6 | [24 empty] | BP25-USN6 | 59580-59586 |  | 0 | [25 empty] | USN6-BP21 via BP25 | 2022.0137 |  | 140.8 | [25 empty] | USN6 wrt BP21 via BP25 | 2022.0137 | [14 empty] | 0
| 59580 | 59586 | BP1J | USN7 | USNO | 59580-59586 | BP1J-USN7 | -95.02 | 0.1 | -90.91 | 0.1 | -95.9 | 0.1 | -92.08 | 0.1 | -96.49 |  |  |  | -82.97 | 0.5 | -84 | 0.2 | [8 empty] | BP1J-USN7 | 59580-59586 | 171.4 | 0 |  | -95.02 | 76.38 | -90.91 | 80.49 | -95.9 | 75.5 | -92.08 | 79.32 | -96.49 | 74.91 |  |  | -82.97 | 88.43 | -84 | 87.4 | [8 empty] | USN7-BP21 via BP1J | 2022.0137 |  | 140.8 |  | -15.865 | 124.935 | -19.085 | 121.715 | -15.78 | 125.02 |  |  | -15.785 | 125.015 |  |  | -15.28 | 125.52 | -16.285 | 124.515 | [8 empty] | USN7 wrt BP21 via BP1J | 2022.0137 | 153.335 | 149.015 | 155.62 |  | 155.715 |  | 153.22 | 153.115 |  |  |  |  | PolaRx5 |  | 160.0094
| 59580 | 59586 | BP25 | USN7 | USNO | 59580-59586 | BP25-USN7 | [24 empty] | BP25-USN7 | 59580-59586 | [27 empty] | USN7-BP21 via BP25 | 2022.0137 |  | 140.8 | [25 empty] | USN7 wrt BP21 via BP25 | 2022.0137
| 59580 | 59586 | BP1J | USN8 | USNO | 59580-59586 | BP1J-USN8 | -142.01 | 0.1 | -138.98 | 0.1 | -142.93 | 0.1 | -140.09 | 0.1 | -143.54 | 0.1 | -135.56 | 0.1 | -131.93 | 0.5 | -131.33 | 0.2 | [8 empty] | BP1J-USN8 | 59580-59586 | 171.4 | 0 |  | -142.01 | 29.39 | -138.98 | 32.42 | -142.93 | 28.47 |  |  | -143.54 | 27.86 | -135.56 | 35.84 | -131.93 | 39.47 | -131.33 | 40.07 | [8 empty] | USN8-BP21 via BP1J | 2022.0137 |  | 140.8 |  | 31.125 | 171.925 | 28.985 | 169.785 | 31.25 | 172.05 |  |  | 31.265 | 172.065 | 33.105 | 173.905 | 33.68 | 174.48 | 31.045 | 171.845 | [8 empty] | USN8 wrt BP21 via BP1J | 2022.0137 | 200.325 | 197.085 | 202.65 |  | 202.765 | 204.805 | 202.18 | 200.445 |  |  |  |  | PolaRx5 |  | 205.3308
| 59580 | 59586 | BP25 | USN8 | USNO | 59580-59586 | BP25-USN8 | [24 empty] | BP25-USN8 | 59580-59586 |  | 0 | [25 empty] | USN8-BP21 via BP25 | 2022.0137 |  | 140.8 | [25 empty] | USN8 wrt BP21 via BP25 | 2022.0137 | [14 empty] | 0
| 59580 | 59586 | BP1J | USN9 | USNO | 59580-59586 | BP1J-USN9 | -143.97 | 0.1 | -161.3 | 0.1 | -143.34 | 0.1 | -159.31 | 0.1 | -145.02 | 0.1 | -160.18 | 0.1 | -161.93 | 0.2 | -153.92 | 0.2 | [8 empty] | BP1J-USN9 | 59580-59586 | 171.4 | 0 |  | -143.97 | 27.43 | -161.3 | 10.1 | -143.34 | 28.06 | -159.31 | 12.09 | -145.02 | 26.38 | -160.18 | 11.22 | -161.93 | 9.47 | -153.92 | 17.48 | [8 empty] | USN9-BP21 via BP1J | 2022.0137 |  | 140.8 |  | 33.085 | 173.885 | 51.305 | 192.105 | 31.66 | 172.46 |  |  | 32.745 | 173.545 | 57.725 | 198.525 | 63.68 | 204.48 | 53.635 | 194.435 | [8 empty] | USN9 wrt BP21 via BP1J | 2022.0137 | 202.285 | 219.405 | 203.06 |  | 204.245 | 229.425 | 232.18 | 223.035 |  |  |  |  | Novatel |  | 175.8346
| 59580 | 59586 | BP25 | USN9 | USNO | 59580-59586 | BP25-USN9 | [24 empty] | BP25-USN9 | 59580-59586 |  | 0 | [25 empty] | USN9-BP21 via BP25 | 2022.0137 |  | 140.8 | [25 empty] | USN9 wrt BP21 via BP25 | 2022.0137 | [14 empty] | 0
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [26 empty] | The colums below make up Table 3 (Traveling - Reference)
BP21 reference (Auto-compensation ON to be used)
| Start | End | Travel | Local | Labo | Date | Pair | RDIF(P1) | Unc | RDIF(P2) | Unc | RDIF(C1) | Unc | RDIF(C2) | Unc | RDIF(E1) | Unc | RDIF(E5a) | Unc | RDIF(D1) | Unc | RDIF(D2) | Unc | RDIF(B1) | Unc | RDIF(B2I) | Unc | RDIF(B1C) | Unc | RDIF(B2a) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a |  | D1 |  | D2 |  | B1 |  | B2 |  | BC |  | B5
[37 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [8 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59440 | 59449 | BP1J | BP21 | BIPM | 59440-59449 | BP1J-BP21 | -90.5 | 0.1 | -89.63 | 0.1 | -91.32 | 0.1 | -90.51 | 0.1 | -91.9 | 0.1 | -82.13 | 0.1 | -77.92 | 0.5 | -79.99 | 0.2 | [8 empty] | BP1J-BP21 | 59440-59449 | 194.46 | 43.3 |  | -90.5 | 60.66 | -89.63 | 61.53 | -91.32 | 59.84 | -90.51 | 60.65 | -91.9 | 59.26 | -82.13 | 69.03 | -77.92 | 73.24 | -79.99 | 71.17
| 59663 | 59667 | BP1J | BP21 | BIPM | 59663-59667 | BP1J-BP21 | -90.72 | 0.1 | -89.81 | 0.1 | -91.49 | 0.1 | -90.61 | 0.1 | -92.1 | 0.1 | -82.23 | 0.1 | -78.03 | 0.5 | -80.03 | 0.2 | [8 empty] | BP1J-BP21 | 59663-59667 | 194.38 | 43.29 |  | -90.72 | 60.37 | -89.81 | 61.28 | -91.49 | 59.6 | -90.61 | 60.48 | -92.1 | 58.99 | -82.23 | 68.86 | -78.03 | 73.06 | -80.03 | 71.06
Results for Travel system #1 | [33 empty] | Misclosure |  |  |  | 0.29 |  | 0.25 |  | 0.24 |  | 0.17 |  | 0.27 |  | 0.17 |  | 0.18 |  | 0.11
|  |  | BP1J | BP21 | BIPM | [26 empty] | BP1J-BP21 |  | Mean (chosen value) |  |  |  | 60.515 |  | 61.405 |  | 59.72 |  | 60.565 |  | 59.125 |  | 68.945 |  | 73.15 |  | 71.115
[34 empty] | Mean (arithmetic) |  |  |  | 60.515 |  | 61.405 |  | 59.72 |  | 60.565 |  | 59.125 |  | 68.945 |  | 73.15 |  | 71.115
| 59440 | 59449 | BP25 | BP21 | BIPM | 59440-59449 | BP25-BP21 | 12.12 | 0.1 | 17.63 | 0.1 | 10.75 | 0.1 | 16.15 | 0.1 | 10.35 | 0.1 | 16.3 | 0.1 | 14.22 | 0.2 | 19.61 | 0.2 | 18.85 | 0.2 | 19.09 | 0.2 | 10.04 | 0.1 | 15.9 | 0.1 | BP25-BP21 | 59440-59449 | 53.49 | 43.3 |  | 12.12 | 22.31 | 17.63 | 27.82 | 10.75 | 20.94 | 16.15 | 26.34 | 10.35 | 20.54 | 16.3 | 26.49 | 14.22 | 24.41 | 19.61 | 29.8 | 18.85 | 29.04 | 19.09 | 29.28 | 10.04 | 20.23 | 15.9 | 26.09
| 59663 | 59667 | BP25 | BP21 | BIPM | 59663-59667 | BP25-BP21 | 12 | 0.1 | 17.61 | 0.1 | 10.66 | 0.1 | 16.13 | 0.1 | 10.27 | 0.1 | 16.41 | 0.1 | 14.02 | 0.2 | 19.53 | 0.2 | 18.73 | 0.2 | 18.85 | 0.2 | 9.95 | 0.1 | 16.05 | 0.1 | BP25-BP21 | 59663-59667 | 53.35 | 43.29 |  | 12 | 22.06 | 17.61 | 27.67 | 10.66 | 20.72 | 16.13 | 26.19 | 10.27 | 20.33 | 16.41 | 26.47 | 14.02 | 24.08 | 19.53 | 29.59 | 18.73 | 28.79 | 18.85 | 28.91 | 9.95 | 20.01 | 16.05 | 26.11
Results for Travel system #2 | [33 empty] | Misclosure |  |  |  | 0.25 |  | 0.15 |  | 0.22 |  | 0.15 |  | 0.21 |  | 0.02 |  | 0.33 |  | 0.21 |  | 0.25 |  | 0.37 |  | 0.22 |  | 0.02
|  |  | BP25 | BP21 | BIPM | [26 empty] | BP25-BP21 |  | Mean (chosen value) |  |  |  | 22.185 |  | 27.745 |  | 20.83 |  | 26.265 |  | 20.435 |  | 26.48 |  | 24.245 |  | 29.695 |  | 28.915 |  | 29.095 |  | 20.12 |  | 26.1
[34 empty] | Mean (arithmetic) |  |  |  | 22.185 |  | 27.745 |  | 20.83 |  | 26.265 |  | 20.435 |  | 26.48 |  | 24.245 |  | 29.695 |  | 28.915 |  | 29.095 |  | 20.12 |  | 26.1
BP1K check
| Start | End | Travel | Local | Labo | Date | Pair | RDIF(P1) | Unc | RDIF(P2) | Unc | RDIF(C1) | Unc | RDIF(C2) | Unc | RDIF(E1) | Unc | RDIF(E5a) | Unc | RDIF(D1) | Unc | RDIF(D2) | Unc | RDIF(B1) | Unc | RDIF(B2I) | Unc | RDIF(B1C) | Unc | RDIF(B2a) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a
[37 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59440 | 59449 | BP1K | BP21 | BIPM | 59440-59449 | BP1K-BP21 | 7.94 | 0.1 | 9.3 | 99 | 7.31 | 0.1 | 8.76 | 0.1 | 7.57 | 0.1 | 5.18 | 0.1 | [12 empty] | BP1K-BP21 | 59440-59449 | -31.26 | 43.3 |  | 7.94 | -66.62 | 9.3 |  | 7.31 | -67.25 | 8.76 | -65.8 | 7.57 | -66.99 | 5.18 | -69.38
| 59663 | 59667 | BP1K | BP21 | BIPM | 59663-59667 | BP1K-BP21 | 7.68 | 0.1 | 8.89 | 0.1 | 7.1 | 0.1 | 8.42 | 0.1 | 7.37 | 0.1 | 4.86 | 0.1 | [12 empty] | BP1K-BP21 | 59663-59667 | -30.66 | 43.29 |  | 7.68 | -66.27 | 8.89 | -65.06 | 7.1 | -66.85 | 8.42 | -65.53 | 7.37 | -66.58 | 4.86 | -69.09
Results for BP1K | [33 empty] | Misclosure |  |  |  | 0.35 |  | 0 |  | 0.4 |  | 0.27 |  | 0.41 |  | 0.29
|  |  | BP1K | BP21 | BIPM | [26 empty] | BP1K-BP21 |  | Mean (chosen value) |  |  |  | -66.445 |  | -65.06 |  | -67.05 |  | -65.665 |  | -66.785 |  | -69.235
[34 empty] | Mean (arithmetic) |  |  |  | -66.445 |  | -65.06 |  | -67.05 |  | -65.665 |  | -66.785 |  | -69.235
BP27 check
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | [18 empty] | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1
| 59440 | 59449 | BP27 | BP21 | BIPM | 59440-59449 | BP27-BP21 | 42.59 | 0.1 | 41.59 | 0.1 | 42.82 | 0.1 | 41.59 | 0.1 | 42.63 | 0.1 | 44.17 | 0.1 | 41.05 | 0.1 | 41.07 | 0.1 | 43.1 | 0.1 | 42.13 | 0.1 | 42.75 | 0.1 | 44.15 | 0.1 | BP27-BP21 | 59440-59449 | 61.34 | 43.25 |  | 42.59 | 60.68 | 41.59 | 59.68 | 42.82 | 60.91 | 41.59 | 59.68 | 42.63 | 60.72 | 44.17 | 62.26 | 41.05 | 59.14 | 41.07 | 59.16 | 43.1 | 61.19 | 42.13 | 60.22 | 42.75 | 60.84 | 44.15 | 62.24
| 59663 | 59667 | BP27 | BP21 | BIPM | 59663-59667 | BP27-BP21 | 75.54 | 0.1 | 74.48 | 0.1 | 75.8 | 0.1 | 74.52 | 0.1 | 75.63 | 0.1 | 77.13 | 0.1 | 73.97 | 0.1 | 73.95 | 0.1 | 76.03 | 0.1 | 74.76 | 0.1 | 75.75 | 0.1 | 77.13 | 0.1 | BP27-BP21 | 59663-59667 | 28.2 | 43.29 |  | 75.54 | 60.45 | 74.48 | 59.39 | 75.8 | 60.71 | 74.52 | 59.43 | 75.63 | 60.54 | 77.13 | 62.04 | 73.97 | 58.88 | 73.95 | 58.86 | 76.03 | 60.94 | 74.76 | 59.67 | 75.75 | 60.66 | 77.13 | 62.04
Results for BP27 | [33 empty] | Misclosure |  |  |  | 0.23 |  | 0.29 |  | 0.2 |  | 0.25 |  | 0.18 |  | 0.22 |  | 0.26 |  | 0.3 |  | 0.25 |  | 0.55 |  | 0.18 |  | 0.2
|  |  | BP27 | BP21 | BIPM | [26 empty] | BP27-BP21 |  | Mean (chosen value) |  |  |  | 60.565 |  | 59.535 |  | 60.81 |  | 59.555 |  | 60.63 |  | 62.15 |  | 59.01 |  | 59.01 |  | 61.065 |  | 59.945 |  | 60.75 |  | 62.14
[34 empty] | Mean (arithmetic) |  |  |  | 60.565 |  | 59.535 |  | 60.81 |  | 59.555 |  | 60.63 |  | 62.15 |  | 59.01 |  | 59.01 |  | 61.065 |  | 59.945 |  | 60.75 |  | 62.14

## Phase 2
Phase 2 of the calibration trip is BIPM-ROA-PTB-OP-BIPM and two traveling systems.
Updated 23/02/2021 for BIPM start; 19/03 for ROA; 29/03 and 01/04 for PTB (preliminary); 28/04 for OP (preliminary); 12/05 for GLO and BDS results for all possible; 09/06: Closure at BIPM and finalization OP.
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [21 empty] | Reference delays measured in the actual set-up. | [24 empty] | Measured / known antenna cable delays
[30 empty] | The column "Note" is present to match the Table in the report | [24 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed.
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [22 empty] | The colums below make up Table 4 (Traveling - Visited) | [24 empty] | The colums below make up Table 5 (Visited - Reference) | [22 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
[76 empty] | BP21 reference values 1001-2020 |  | 28.4 | 27.3 | 30.6 | 0 | 30.7 | 30.9 | 27.7 | 28.6 | 24.4 | 25.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a |  | D1 |  | D2 |  | B1 |  | B2 |  | Pair | Date | CABDLYV | CABDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | E1 |  | E5 |  | D1 |  | D2 |  | B1 |  | B2 |  | Pair | Date | P1 (ns) | P2 (ns) | C1 (ns) | C2 (ns) | E1 (ns) | E5 (ns) | D1(ns) | D2(ns) | B1(ns) | B2(ns)
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY |  |  | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | INTDLYV | [7 empty] | Type | (CAB-REF)apriori | TOTDLY(P3)
First set-up with antenna cables of the two traveling receivers interchanged
| 59299 | 59301 | BP1J | PTBB | PTB | 59299-59301 | BP1J-PTBB | -161.36 | 0.1 | -159.11 | 0.1 | -162.02 | 0.1 | -160.15 | 0.1 | -162.54 | 0.1 | -152.65 | 0.1 | [8 empty] | BP1J-PTBB | 59299-59301 | 208.64 | 54.3 |  | -161.36 | -7.02 | -159.11 | -4.77 | -162.02 | -7.68 | -160.15 | -5.81 | -162.54 | -8.2 | -152.65 | 1.69 | [8 empty] | PTBB-BP21 via BP1J | 2021.2384 | 205.7 | 140.8 |  | 67.25 | 2.35 | 65.96 | 1.06 | 67.105 | 2.205 | 67.01 | 2.11 | 66.88 | 1.98 | [8 empty] | PTBB wrt BP21 via BP1J | 2021.2384 | 30.75 | 28.36 | 32.805 |  | 32.81 | 32.88 |  |  |  |  | PolaRx5 |  | 185.8426
| 59299 | 59301 | BP25 | PTBB | PTB | 59299-59301 | BP25-PTBB | -57.45 | 0.1 | -50.86 | 0.1 | -58.62 | 0.1 | -52.42 | 0.1 | -59.08 | 0.1 | -52.77 | 0.1 | [8 empty] | BP25-PTBB | 59299-59301 | 68.39 | 54.3 |  | -57.45 | -43.36 | -50.86 | -36.77 | -58.62 | -44.53 | -52.42 | -38.33 | -59.08 | -44.99 | -52.77 | -38.68 | [8 empty] | PTBB-BP21 via BP25 | 2021.2384 | 205.7 | 140.8 |  | 65.455 | 0.555 | 64.32 | -0.58 | 65.215 | 0.315 | 65.28 | 0.38 | 65.055 | 0.155 | [8 empty] | PTBB wrt BP21 via BP25 | 2021.2384 | 28.955 | 26.72 | 30.915 |  | 31.08 | 31.055 |  |  |  |  | PolaRx5 |  | 183.8081
| 59299 | 59301 | BP1J | PT07 | PTB | 59299-59301 | BP1J-PT07 | 19.75 | 0.1 | 19.74 | 0.1 | 21.2 | 0.1 | [14 empty] | BP1J-PT07 | 59299-59301 | 208.64 | 43.4 | * | 19.75 | 184.99 | 19.74 | 184.98 | 21.2 | 186.44 | [14 empty] | PT07-BP21 via BP1J | 2021.2384 | 245.8 | 140.8 | * | -124.76 | -229.76 | -123.79 | -228.79 | -127.015 | -232.015 | [12 empty] | PT07 wrt BP21 via BP1J | 2021.2384 | 0.94 | 0.81 | 0.885 | [7 empty] | GTR50 | 202.3 | 1.2408
| 59299 | 59301 | BP25 | PT07 | PTB | 59299-59301 | BP25-PT07 | 123.65 | 0.1 | 128 | 0.1 | 124.55 | 0.1 | [14 empty] | BP25-PT07 | 59299-59301 | 68.39 | 43.4 | * | 123.65 | 148.64 | 128 | 152.99 | 124.55 | 149.54 | [14 empty] | PT07-BP21 via BP25 | 2021.2384 | 245.8 | 140.8 | * | -126.545 | -231.545 | -125.44 | -230.44 | -128.855 | -233.855 | [12 empty] | PT07 wrt BP21 via BP25 | 2021.2384 | -0.845 | -0.84 | -0.955 | [7 empty] | GTR50 | 202.3 | -0.75273
| 59299 | 59301 | BP1J | PT09 | PTB | 59299-59301 | BP1J-PT09 | -52.11 | 0.1 | -51.36 | 0.1 | -52.13 | 0.1 | -51.39 | 0.1 | -52.1 | 0.1 | -51.42 | 0.1 | [8 empty] | BP1J-PT09 | 59299-59301 | 208.64 | 183.2 |  | -52.11 | -26.67 | -51.36 | -25.92 | -52.13 | -26.69 | -51.39 | -25.95 | -52.1 | -26.66 | -51.42 | -25.98 | [8 empty] | PT09-BP21 via BP1J | 2021.2384 | 198.7 | 140.8 |  | 86.9 | 29 | 87.11 | 29.21 | 86.115 | 28.215 | 85.47 | 27.57 | 94.55 | 36.65 | [8 empty] | PT09 wrt BP21 via BP1J | 2021.2384 | 57.4 | 56.51 | 58.815 |  | 58.27 | 67.55 |  |  |  |  | PolaRx4 |  | 74.2751
| 59299 | 59301 | BP25 | PT09 | PTB | 59299-59301 | BP25-PT09 | 51.78 | 0.1 | 56.86 | 0.1 | 51.2 | 0.1 | 56.37 | 0.1 | 51.33 | 0.1 | 48.51 | 0.1 | [8 empty] | BP25-PT09 | 59299-59301 | 68.39 | 183.2 |  | 51.78 | -63.03 | 56.86 | -57.95 | 51.2 | -63.61 | 56.37 | -58.44 | 51.33 | -63.48 | 48.51 | -66.3 | [8 empty] | PT09-BP21 via BP25 | 2021.2384 | 198.7 | 140.8 |  | 85.125 | 27.225 | 85.5 | 27.6 | 84.295 | 26.395 | 83.77 | 25.87 | 92.675 | 34.775 | [8 empty] | PT09 wrt BP21 via BP25 | 2021.2384 | 55.625 | 54.9 | 56.995 |  | 56.57 | 65.675 |  |  |  |  | PolaRx4 |  | 72.2451
| 59299 | 59301 | BP1J | PT10 | PTB | 59299-59301 | BP1J-PT10 | -225.13 | 0.1 | -218.69 | 0.1 | -225.58 | 0.1 | -219.41 | 0.1 | -227.51 | 0.1 | -206.64 | 0.1 | [8 empty] | BP1J-PT10 | 59299-59301 | 208.64 | 36.8 |  | -225.13 | -53.29 | -218.69 | -46.85 | -225.58 | -53.74 | -219.41 | -47.57 | -227.51 | -55.67 | -206.64 | -34.8 | [8 empty] | PT10-BP21 via BP1J | 2021.2384 | 250 | 140.8 |  | 113.52 | 4.32 | 108.04 | -1.16 | 113.165 | 3.965 | 114.48 | 5.28 | 103.37 | -5.83 | [8 empty] | PT10 wrt BP21 via BP1J | 2021.2384 | 32.72 | 26.14 | 34.565 |  | 35.98 | 25.07 |  |  |  |  | GTR51 | 0 | 256.0861
| 59299 | 59301 | BP25 | PT10 | PTB | 59299-59301 | BP25-PT10 | -121.27 | 0.1 | -110.4 | 0.1 | -122.27 | 0.1 | -111.59 | 0.1 | -124.08 | 0.1 | -106.78 | 0.1 | [8 empty] | BP25-PT10 | 59299-59301 | 68.39 | 36.8 |  | -121.27 | -89.68 | -110.4 | -78.81 | -122.27 | -90.68 | -111.59 | -80 | -124.08 | -92.49 | -106.78 | -75.19 | [8 empty] | PT10-BP21 via BP25 | 2021.2384 | 250 | 140.8 |  | 111.775 | 2.575 | 106.36 | -2.84 | 111.365 | 2.165 | 112.78 | 3.58 | 101.565 | -7.635 | [8 empty] | PT10 wrt BP21 via BP25 | 2021.2384 | 30.975 | 24.46 | 32.765 |  | 34.28 | 23.265 |  |  |  |  | GTR51 | 0 | 254.2407
* PT07 is a GTR50 with REFDLY=43.5 ns and CABDLY=245.8 ns introduced a priori
Second set-up with the antenna cables properly attached to the correct receivers
| 59303 | 59311 | BP1J | PTBB | PTB | 59303-59311 | BP1J-PTBB | -160.53 | 0.1 | -158.32 | 0.1 | -161.18 | 0.1 | -159.34 | 0.1 | -161.69 | 0.1 | -151.69 | 0.1 | -147.81 | 0.5 | -150.54 | 0.2 |  |  |  |  | BP1J-PTBB | 59303-59311 | 208.91 | 56.22 |  | -160.53 | -7.84 | -158.32 | -5.63 | -161.18 | -8.49 | -159.34 | -6.65 | -161.69 | -9 | -151.69 | 1 | -147.81 | 4.88 | -150.54 | 2.15 |  |  |  |  | PTBB-BP21 via BP1J | 2021.2575 | 205.7 | 140.8 |  | 68.07 | 3.17 | 66.82 | 1.92 | 67.915 | 3.015 | 67.81 | 2.91 | 67.57 | 2.67 | 67.885 | 2.985 | 68.58 | 3.68 |  |  |  |  | PTBB wrt BP21 via BP1J | 2021.2575 | 31.57 | 29.22 | 33.615 |  | 33.61 | 33.57 | 30.685 | 32.28 |  |  | PolaRx5 |  | 184.6808
| 59303 | 59311 | BP25 | PTBB | PTB | 59303-59311 | BP25-PTBB | -58.24 | 0.1 | -51.6 | 0.1 | -59.4 | 0.1 | -53.14 | 0.1 | -59.84 | 0.1 | -53.46 | 0.1 | -56.57 | 0.2 | -51.52 | 0.1 | -57.6 | 0.1 | -55.72 | 5 | BP25-PTBB | 59303-59311 | 68.48 | 56.22 |  | -58.24 | -45.98 | -51.6 | -39.34 | -59.4 | -47.14 | -53.14 | -40.88 | -59.84 | -47.58 | -53.46 | -41.2 | -56.57 | -44.31 | -51.52 | -39.26 | -57.6 | -45.34 | -55.72 | -43.46 | PTBB-BP21 via BP25 | 2021.2575 | 205.7 | 140.8 |  | 68.075 | 3.175 | 66.89 | 1.99 | 67.825 | 2.925 | 67.87 | 2.97 | 67.575 | 2.675 | 68.35 | 3.45 | 68.8 | 3.9 | 74.125 | 9.225 | 72.7 | 7.8 | PTBB wrt BP21 via BP25 | 2021.2575 | 31.575 | 29.29 | 33.525 |  | 33.67 | 33.575 | 31.15 | 32.5 | 33.625 | 33.3 | PolaRx5 |  | 184.5853
| 59303 | 59311 | BP1J | PT07 | PTB | 59303-59311 | BP1J-PT07 | 20.66 | 0.1 | 20.55 | 0.1 | 22.14 | 0.1 | [14 empty] | BP1J-PT07 | 59303-59311 | 208.91 | 43.36 | * | 20.66 | 186.21 | 20.55 | 186.1 | 22.14 | 187.69 | [14 empty] | PT07-BP21 via BP1J | 2021.2575 | 245.8 | 140.8 | * | -125.98 | -230.98 | -124.91 | -229.91 | -128.265 | -233.265 | [12 empty] | PT07 wrt BP21 via BP1J | 2021.2575 | -0.28 | -0.31 | -0.365 | [7 empty] | GTR50 | 202.3 | -0.09365
| 59303 | 59311 | BP25 | PT07 | PTB | 59303-59311 | BP25-PT07 | 122.91 | 0.1 | 127.3 | 0.1 | 123.85 | 0.1 | [14 empty] | BP25-PT07 | 59303-59311 | 68.48 | 43.36 | * | 122.91 | 148.03 | 127.3 | 152.42 | 123.85 | 148.97 | [14 empty] | PT07-BP21 via BP25 | 2021.2575 | 245.8 | 140.8 | * | -125.935 | -230.935 | -124.87 | -229.87 | -128.285 | -233.285 | [12 empty] | PT07 wrt BP21 via BP25 | 2021.2575 | -0.235 | -0.27 | -0.385 | [7 empty] | GTR50 | 202.3 | -0.040925
| 59303 | 59311 | BP1J | PT09 | PTB | 59303-59311 | BP1J-PT09 | -51.25 | 0.1 | -50.62 | 0.1 | -51.25 | 0.1 | -50.62 | 0.1 | -51.26 | 0.1 | -50.36 | 0.1 | -51.11 | 0.1 | -49.12 | 0.1 |  |  |  |  | BP1J-PT09 | 59303-59311 | 208.91 | 182.88 |  | -51.25 | -25.22 | -50.62 | -24.59 | -51.25 | -25.22 | -50.62 | -24.59 | -51.26 | -25.23 | -50.36 | -24.33 | [8 empty] | PT09-BP21 via BP1J | 2021.2575 | 198.7 | 140.8 |  | 85.45 | 27.55 | 85.78 | 27.88 | 84.645 | 26.745 | 84.04 | 26.14 | 92.9 | 35 | [8 empty] | PT09 wrt BP21 via BP1J | 2021.2575 | 55.95 | 55.18 | 57.345 |  | 56.84 | 65.9 |  |  |  |  | PolaRx4 |  | 72.9597
| 59303 | 59311 | BP25 | PT09 | PTB | 59303-59311 | BP25-PT09 | 51.01 | 0.1 | 56.16 | 0.1 | 50.44 | 0.1 | 55.64 | 0.1 | 50.59 | 0.1 | 47.82 | 0.1 | 40.28 | 0.5 | 40.71 | 0.1 |  |  | 49.46 | 5 | BP25-PT09 | 59303-59311 | 68.48 | 182.88 |  | 51.01 | -63.39 | 56.16 | -58.24 | 50.44 | -63.96 | 55.64 | -58.76 | 50.59 | -63.81 | 47.82 | -66.58 | [8 empty] | PT09-BP21 via BP25 | 2021.2575 | 198.7 | 140.8 |  | 85.485 | 27.585 | 85.79 | 27.89 | 84.645 | 26.745 | 84.1 | 26.2 | 92.955 | 35.055 | [8 empty] | PT09 wrt BP21 via BP25 | 2021.2575 | 55.985 | 55.19 | 57.345 |  | 56.9 | 65.955 |  |  |  |  | PolaRx4 |  | 73.0333
| 59303 | 59311 | BP1J | PT10 | PTB | 59303-59311 | BP1J-PT10 | -224.26 | 0.1 | -217.89 | 0.1 | -224.71 | 0.1 | -218.6 | 0.1 | -226.64 | 0.1 | -205.65 | 0.1 | -212.02 | 0.5 | -217.8 | 0.2 |  |  |  |  | BP1J-PT10 | 59303-59311 | 208.91 | 36.56 |  | -224.26 | -51.91 | -217.89 | -45.54 | -224.71 | -52.36 | -218.6 | -46.25 | -226.64 | -54.29 | -205.65 | -33.3 | [8 empty] | PT10-BP21 via BP1J | 2021.2575 | 250 | 140.8 |  | 112.14 | 2.94 | 106.73 | -2.47 | 111.785 | 2.585 | 113.1 | 3.9 | 101.87 | -7.33 | [8 empty] | PT10 wrt BP21 via BP1J | 2021.2575 | 31.34 | 24.83 | 33.185 |  | 34.6 | 23.57 |  |  |  |  | GTR51 | 0 | 254.8379
| 59303 | 59311 | BP25 | PT10 | PTB | 59303-59311 | BP25-PT10 | -122.01 | 0.1 | -111.11 | 0.1 | -123.02 | 0.1 | -112.34 | 0.1 | -124.83 | 0.1 | -107.45 | 0.1 | -120.75 | 0.2 | -119 | 0.1 |  |  |  |  | BP25-PT10 | 59303-59311 | 68.48 | 36.56 |  | -122.01 | -90.09 | -111.11 | -79.19 | -123.02 | -91.1 | -112.34 | -80.42 | -124.83 | -92.91 | -107.45 | -75.53 | [8 empty] | PT10-BP21 via BP25 | 2021.2575 | 250 | 140.8 |  | 112.185 | 2.985 | 106.74 | -2.46 | 111.785 | 2.585 | 113.2 | 4 | 101.905 | -7.295 | [8 empty] | PT10 wrt BP21 via BP25 | 2021.2575 | 31.385 | 24.84 | 33.185 |  | 34.7 | 23.605 |  |  |  |  | GTR51 | 0 | 254.937
* PT07 is a GTR50 with REFDLY=43.5 ns and CABDLY=245.8 ns introduced a priori
| 59281 | 59287 | BP1J | RO_5 | ROA | 59281-59287 | BP1J-RO_5 | -234.52 | 0.1 | -234.88 | 0.1 | -232.5 | 0.1 | [14 empty] | BP1J-RO_5 | 59281-59287 | 464.5 | 306.8 | ** | -234.52 | -76.82 | -234.88 | -77.18 | -232.5 | -74.8 | [14 empty] | RO_5-BP21 via BP1J | 2021.1945 | 91.5 | 140.8 | ** | 137.05 | 186.35 | 138.37 | 187.67 | 134.225 | 183.525 | [12 empty] | RO_5 wrt BP21 via BP1J | 2021.1945 | -0.35 | -0.13 | -0.975 | [7 empty] | GTR50 | -215.1 | -0.8899
| 59281 | 59287 | BP25 | RO_5 | ROA | 59281-59287 | BP25-RO_5 | -139.06 | 0.1 | -134.86 | 0.2 | -137.51 | 0.1 | [14 empty] | BP25-RO_5 | 59281-59287 | 330.92 | 306.8 | ** | -139.06 | -114.94 | -134.86 | -110.74 | -137.51 | -113.39 | [14 empty] | RO_5-BP21 via BP25 | 2021.1945 | 91.5 | 140.8 | ** | 137.035 | 186.335 | 138.29 | 187.59 | 134.075 | 183.375 | [12 empty] | RO_5 wrt BP21 via BP25 | 2021.1945 | -0.365 | -0.21 | -1.125 | [7 empty] | GTR50 | -215.1 | -0.80448
| 59281 | 59287 | BP1J | RO_6 | ROA | 59281-59287 | BP1J-RO_6 | 113.69 | 0.1 | 114.58 | 0.1 | 113.49 | 0.1 | 115.08 | 0.1 |  |  |  |  | 113.78 | 1 | 125.62 | 0.5 |  |  |  |  | BP1J-RO_6 | 59281-59287 | 464.5 | 485.1 |  | 113.69 | 93.09 | 114.58 | 93.98 | 113.49 | 92.89 | 115.08 | 94.48 | [12 empty] | RO_6-BP21 via BP1J | 2021.1945 | 82 | 140.8 |  | -32.86 | 25.94 | -32.79 | 26.01 | -33.465 | 25.335 | [12 empty] | RO_6 wrt BP21 via BP1J | 2021.1945 | 54.34 | 53.31 | 55.935 | [7 empty] | PolaRx3 |  | -347.1687
| 59281 | 59287 | BP25 | RO_6 | ROA | 59281-59287 | BP25-RO_6 | 209.2 | 0.1 | 214.66 | 0.1 | 208.51 | 0.1 | 214.69 | 0.1 |  |  |  |  | 199.39 | 3 | 218 | 0.5 |  |  |  |  | BP25-RO_6 | 59281-59287 | 330.92 | 485.1 |  | 209.2 | 55.02 | 214.66 | 60.48 | 208.51 | 54.33 | 214.69 | 60.51 | [12 empty] | RO_6-BP21 via BP25 | 2021.1945 | 82 | 140.8 |  | -32.925 | 25.875 | -32.93 | 25.87 | -33.645 | 25.155 | [12 empty] | RO_6 wrt BP21 via BP25 | 2021.1945 | 54.275 | 53.17 | 55.755 | [7 empty] | PolaRx3 |  | -347.1178
| 59281 | 59287 | BP1J | RO_7 | ROA | 59281-59287 | BP1J-RO_7 | 73.12 | 0.1 | 73.34 | 0.1 | 73.22 | 0.1 | 73.48 | 0.1 | 73.21 |  | 74.58 |  | 72.13 | 0.1 | 76.46 | 0.1 |  |  |  |  | BP1J-RO_7 | 59281-59287 | 464.5 | 452.4 |  | 73.12 | 85.22 | 73.34 | 85.44 | 73.22 | 85.32 | 73.48 | 85.58 | 73.21 | 85.31 | 74.58 | 86.68 | 72.13 | 84.23 | 76.46 | 88.56 |  |  |  |  | RO_7-BP21 via BP1J | 2021.1945 | 89.9 | 140.8 |  | -24.99 | 25.91 | -24.25 | 26.65 | -25.895 | 25.005 | -26.5 | 24.4 | -18.11 | 32.79 | -11.465 | 39.435 | -17.83 | 33.07 |  |  |  |  | RO_7 wrt BP21 via BP1J | 2021.1945 | 54.31 | 53.95 | 55.605 |  | 55.1 | 63.69 | 67.135 | 61.67 |  |  | PolaRx4 |  | -307.6338
| 59281 | 59287 | BP25 | RO_7 | ROA | 59281-59287 | BP25-RO_7 | 168.62 | 0.1 | 173.44 | 0.1 | 168.21 | 0.1 | 173.1 | 0.1 | 168.29 |  | 165.85 |  | 157.57 | 3 | 168.74 | 1 | 169.6 | 0.2 | 174.09 | 0.2 | BP25-RO_7 | 59281-59287 | 330.92 | 452.4 |  | 168.62 | 47.14 | 173.44 | 51.96 | 168.21 | 46.73 | 173.1 | 51.62 | 168.29 | 46.81 | 165.85 | 44.37 | 157.57 | 36.09 | 168.74 | 47.26 | 169.6 | 48.12 | 174.09 | 52.61 | RO_7-BP21 via BP25 | 2021.1945 | 89.9 | 140.8 |  | -25.045 | 25.855 | -24.41 | 26.49 | -26.045 | 24.855 | -26.52 | 24.38 | -17.995 | 32.905 | -12.05 | 38.85 | -17.72 | 33.18 | -19.335 | 31.565 | -23.37 | 27.53 | RO_7 wrt BP21 via BP25 | 2021.1945 | 54.255 | 53.79 | 55.455 |  | 55.08 | 63.805 | 66.55 | 61.78 | 55.965 | 53.03 | PolaRx4 |  | -307.5266
| 59281 | 59287 | BP1J | RO_8 | ROA | 59281-59287 | BP1J-RO_8 | -441.01 | 0.1 | -439.54 | 0.1 | -441.52 | 0.1 | -440.2 | 0.1 | -443.37 |  | -425.05 |  | -433.85 | 2 | -439.17 | 0.5 |  |  |  |  | BP1J-RO_8 | 59281-59287 | 464.5 | 20.4 |  | -441.01 | 3.09 | -439.54 | 4.56 | -441.52 | 2.58 | -440.2 | 3.9 | -443.37 | 0.73 | -425.05 | 19.05 | [8 empty] | RO_8-BP21 via BP1J | 2021.1945 | 197.1 | 140.8 |  | 57.14 | 0.84 | 56.63 | 0.33 | 56.845 | 0.545 | 58.08 | 1.78 | 49.52 | -6.78 | [8 empty] | RO_8 wrt BP21 via BP1J | 2021.1945 | 29.24 | 27.63 | 31.145 |  | 32.48 | 24.12 |  |  |  |  | GTR51 | 0 | 208.4274
| 59281 | 59287 | BP25 | RO_8 | ROA | 59281-59287 | BP25-RO_8 | -345.47 | 0.1 | -339.43 | 0.1 | -346.55 | 0.1 | -340.57 | 0.1 | -348.29 |  | -333.81 |  | -349 | 2 | -346.97 | 0.5 |  |  |  |  | BP25-RO_8 | 59281-59287 | 330.92 | 20.4 |  | -345.47 | -34.95 | -339.43 | -28.91 | -346.55 | -36.03 | -340.57 | -30.05 | -348.29 | -37.77 | -333.81 | -23.29 | [8 empty] | RO_8-BP21 via BP25 | 2021.1945 | 197.1 | 140.8 |  | 57.045 | 0.745 | 56.46 | 0.16 | 56.715 | 0.415 | 58.06 | 1.76 | 49.665 | -6.635 | [8 empty] | RO_8 wrt BP21 via BP25 | 2021.1945 | 29.145 | 27.46 | 31.015 |  | 32.46 | 24.265 |  |  |  |  | GTR51 | 0 | 208.4483
| 59281 | 59287 | BP1J | RO_9 | ROA | 59281-59287 | BP1J-RO_9 | 102.12 | 0.1 | 102.83 | 0.1 | 102.09 | 0.1 | 102.9 | 0.1 | 102.11 |  | 103.26 |  | 102.75 | 0.1 | 104.89 | 0.1 |  |  |  |  | BP1J-RO_9 | 59281-59287 | 464.5 | 451.8 |  | 102.12 | 114.82 | 102.83 | 115.53 | 102.09 | 114.79 | 102.9 | 115.6 | 102.11 | 114.81 | 103.26 | 115.96 | 102.75 | 115.45 | 104.89 | 117.59 |  |  |  |  | RO_9-BP21 via BP1J | 2021.1945 | 59.7 | 140.8 |  | -54.59 | 26.51 | -54.34 | 26.76 | -55.365 | 25.735 | -56 | 25.1 | -47.39 | 33.71 | -42.685 | 38.415 | -46.86 | 34.24 |  |  |  |  | RO_9 wrt BP21 via BP1J | 2021.1945 | 54.91 | 54.06 | 56.335 |  | 55.8 | 64.61 | 66.115 | 62.84 |  |  | PolaRx4 |  | -335.8767
| 59281 | 59287 | BP25 | RO_9 | ROA | 59281-59287 | BP25-RO_9 | 197.61 | 0.1 | 202.92 | 0.1 | 197.06 | 0.1 | 202.52 | 0.1 | 197.19 |  | 194.52 |  | 187.93 | 3 | 197.11 | 0.5 | 198.44 | 0.2 | 202.45 | 0.2 | BP25-RO_9 | 59281-59287 | 330.92 | 451.8 |  | 197.61 | 76.73 | 202.92 | 82.04 | 197.06 | 76.18 | 202.52 | 81.64 | 197.19 | 76.31 | 194.52 | 73.64 | 187.93 | 67.05 | 197.11 | 76.23 | 198.44 | 77.56 | 202.45 | 81.57 | RO_9-BP21 via BP25 | 2021.1945 | 59.7 | 140.8 |  | -54.635 | 26.465 | -54.49 | 26.61 | -55.495 | 25.605 | -56.02 | 25.08 | -47.265 | 33.835 | -43.01 | 38.09 | -46.69 | 34.41 | -48.775 | 32.325 | -52.33 | 28.77 | RO_9 wrt BP21 via BP25 | 2021.1945 | 54.865 | 53.91 | 56.205 |  | 55.78 | 64.735 | 65.79 | 63.01 | 56.725 | 54.27 | PolaRx4 |  | -335.7595
| 59281 | 59287 | BP1J | RO10 | ROA | 59281-59287 | BP1J-RO10 | -457.2 | 0.1 | -454.91 | 0.1 | -457.59 | 0.1 | -455.97 |  | -458.17 |  | -447.77 |  | -444.63 | 3 | -447.26 | 2 |  |  |  |  | BP1J-RO10 | 59281-59287 | 464.5 | 5.1 |  | -457.2 | 2.2 | -454.91 | 4.49 | -457.59 | 1.81 | -455.97 | 3.43 | -458.17 | 1.23 | -447.77 | 11.63 | -444.63 | 14.77 | -447.26 | 12.14 |  |  |  |  | RO10-BP21 via BP1J | 2021.1945 | 199 | 140.8 |  | 58.03 | -0.17 | 56.7 | -1.5 | 57.615 | -0.585 | 57.58 | -0.62 | 56.94 | -1.26 | 57.995 | -0.205 | 58.59 | 0.39 |  |  |  |  | RO10 wrt BP21 via BP1J | 2021.1945 | 28.23 | 25.8 | 30.015 |  | 30.08 | 29.64 | 27.495 | 28.99 |  |  | PolaRx5 |  | 225.8843
| 59281 | 59287 | BP25 | RO10 | ROA | 59281-59287 | BP25-RO10 | -361.73 | 0.1 | -354.83 | 0.1 | -362.58 | 0.1 | -356.39 |  | -363.11 |  | -356.52 |  | -359.82 | 2 | -354.79 | 0.5 | -352.65 | 0.2 | -355.69 | 0.2 | BP25-RO10 | 59281-59287 | 330.92 | 5.1 |  | -361.73 | -35.91 | -354.83 | -29.01 | -362.58 | -36.76 | -356.39 | -30.57 | -363.11 | -37.29 | -356.52 | -30.7 | -359.82 | -34 | -354.79 | -28.97 | -352.65 | -26.83 | -355.69 | -29.87 | RO10-BP21 via BP25 | 2021.1945 | 199 | 140.8 |  | 58.005 | -0.195 | 56.56 | -1.64 | 57.445 | -0.755 | 57.58 | -0.62 | 57.075 | -1.125 | 58.04 | -0.16 | 58.51 | 0.31 | 55.615 | -2.585 | 59.11 | 0.91 | RO10 wrt BP21 via BP25 | 2021.1945 | 28.205 | 25.66 | 29.845 |  | 30.08 | 29.775 | 27.54 | 28.91 | 21.815 | 26.41 | PolaRx5 |  | 226.037
** RO_5 is a GTR50 with REFDLY=306.6 ns and CABDLY=91.5 ns introduced a priori | [52 empty] | * Value in CGGTTS
| 59326 | 59336 | BP1J | OP71 | OP | 59326-59336 | BP1J-OP71 | -13.32 | 0.1 | -12.43 | 0.1 | -13.38 | 0.1 | -12.4 | 0.1 | -13.36 | 0.1 | -12.4 | 0.1 | -13.19 | 0.1 | -11.11 | 0.1 |  |  |  |  | BP1J-OP71 | 59326-59336 | 251.18 | 192.1 |  | -13.32 | 45.76 | -12.43 | 46.65 | -13.38 | 45.7 | -12.4 | 46.68 | -13.36 | 45.72 | -12.4 | 46.68 | -13.19 | 45.89 | -11.11 | 47.97 |  |  |  |  | OP71-BP21 via BP1J | 2021.3233 | 128.7 | 140.8 | * | 14.47 | 26.57 | 14.54 | 26.64 | 13.725 | 25.825 | 13.09 | 25.19 | 21.89 | 33.99 | 26.875 | 38.975 | 22.76 | 34.86 |  |  |  |  | OP71 wrt BP21 via BP1J | 2021.3233 | 54.97 | 53.94 | 56.425 |  | 55.89 | 64.89 | 66.675 | 63.46 |  |  | PolaRx4 |  | -6.8387
| 59326 | 59336 | BP25 | OP71 | OP | 59326-59336 | BP25-OP71 | 94.56 | 0.1 | 100.03 | 0.1 | 93.94 | 0.1 | 99.52 | 0.1 | 94.1 | 0.1 | 91.2 | 0.1 | 84 | 0.5 | 93.23 | 0.2 | 95.33 | 0.2 | 99.74 | 0.2 | BP25-OP71 | 59326-59336 | 105.37 | 192.1 |  | 94.56 | 7.83 | 100.03 | 13.3 | 93.94 | 7.21 | 99.52 | 12.79 | 94.1 | 7.37 | 91.2 | 4.47 | 84 | -2.73 | 93.23 | 6.5 | 95.33 | 8.6 | 99.74 | 13.01 | OP71-BP21 via BP25 | 2021.3233 | 128.7 | 140.8 | * | 14.265 | 26.365 | 14.25 | 26.35 | 13.475 | 25.575 | 12.92 | 25.02 | 21.905 | 34.005 | 26.77 | 38.87 | 23.04 | 35.14 | 20.185 | 32.285 | 16.23 | 28.33 | OP71 wrt BP21 via BP25 | 2021.3233 | 54.765 | 53.65 | 56.175 |  | 55.72 | 64.905 | 66.57 | 63.74 | 56.685 | 53.83 | PolaRx4 |  | -6.9123
| 59326 | 59336 | BP1J | OP73 | OP | 59326-59336 | BP1J-OP73 | -95.7 | 0.1 | -92.65 | 0.1 | -96.53 | 0.1 | -93.74 | 0.1 | -97.06 | 0.1 | -86.55 | 0.1 | -83.25 | 0.5 | -83.04 | 0.2 |  |  |  |  | BP1J-OP73 | 59326-59336 | 251.18 | 85.22 |  | -95.7 | 70.26 | -92.65 | 73.31 | -96.53 | 69.43 | -93.74 | 72.22 | -97.06 | 68.9 | -86.55 | 79.41 | -83.25 | 82.71 | -83.04 | 82.92 |  |  |  |  | OP73-BP21 via BP1J | 2021.3233 | 129.6 | 140.8 | * | -10.03 | 1.17 | -12.12 | -0.92 | -10.005 | 1.195 | -10.09 | 1.11 | -10.84 | 0.36 | -9.945 | 1.255 | -12.19 | -0.99 |  |  |  |  | OP73 wrt BP21 via BP1J | 2021.3233 | 29.57 | 26.38 | 31.795 |  | 31.81 | 31.26 | 28.955 | 27.61 |  |  | PolaRx5/off |  | 78.8785
| 59326 | 59336 | BP25 | OP73 | OP | 59326-59336 | BP25-OP73 | 12.16 | 0.1 | 19.77 | 0.1 | 10.83 | 0.1 | 18.14 | 0.1 | 10.4 | 0.1 | 17.03 | 0.1 | 13.75 | 0.5 | 21.58 | 0.2 | 19.9 | 0.2 | 20.33 | 0.2 | BP25-OP73 | 59326-59336 | 105.37 | 85.22 |  | 12.16 | 32.31 | 19.77 | 39.92 | 10.83 | 30.98 | 18.14 | 38.29 | 10.4 | 30.55 | 17.03 | 37.18 | 13.75 | 33.9 | 21.58 | 41.73 | 19.9 | 40.05 | 20.33 | 40.48 | OP73-BP21 via BP25 | 2021.3233 | 129.6 | 140.8 | * | -10.215 | 0.985 | -12.37 | -1.17 | -10.295 | 0.905 | -10.26 | 0.94 | -10.805 | 0.395 | -9.86 | 1.34 | -12.19 | -0.99 | -11.265 | -0.065 | -11.24 | -0.04 | OP73 wrt BP21 via BP25 | 2021.3233 | 29.385 | 26.13 | 31.505 |  | 31.64 | 31.295 | 29.04 | 27.61 | 24.335 | 25.46 | PolaRx5/off |  | 78.794
| 59326 | 59336 | BP1J | OPMT | OP | 59326-59336 | BP1J-OPMT | -350.06 | 0.1 | -362.17 | 0.1 | -349.11 | 0.1 | [14 empty] | BP1J-OPMT | 59326-59336 | 251.18 | 137.2 |  | -350.06 | -236.08 | -362.17 | -248.19 | -349.11 | -235.13 | [14 empty] | OPMT-BP21 via BP1J | 2021.3233 | 156.5 | 140.8 | * | 296.31 | 280.61 | 309.38 | 293.68 | 294.555 | 278.855 | [12 empty] | OPMT wrt BP21 via BP1J | 2021.3233 | 309.01 | 320.98 | 309.455 | [7 empty] | Z12T |  | 309.8163
| 59326 | 59336 | BP25 | OPMT | OP | 59326-59336 | BP25-OPMT | -242.22 | 0.1 | -249.75 | 0.1 | -241.77 | 0.1 | [14 empty] | BP25-OPMT | 59326-59336 | 105.37 | 137.2 |  | -242.22 | -274.05 | -249.75 | -281.58 | -241.77 | -273.6 | [14 empty] | OPMT-BP21 via BP25 | 2021.3233 | 156.5 | 140.8 | * | 296.145 | 280.445 | 309.13 | 293.43 | 294.285 | 278.585 | [12 empty] | OPMT wrt BP21 via BP25 | 2021.3233 | 308.845 | 320.73 | 309.185 | [7 empty] | Z12T |  | 309.7827
* Value in CGGTTS
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [22 empty] | The colums below make up Table 3 (Traveling - Reference)
BP21 reference (Auto-compensation ON to be used)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a |  | D1 |  | D2 |  | B1 |  | B2
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59261 | 59266 | BP1J | BP21 | BIPM | 59261-59266 | BP1J-BP21 | -92.12 | 0.1 | -91.25 | 0.1 | -92.92 | 0.1 | -92.07 | 0.1 | -93.56 | 0.1 | -83.83 | 0.1 | -79.5 | 0.5 | -81.6 | 0.2 |  |  |  |  | BP1J-BP21 | 59261-59266 | 195.55 | 43.25 |  | -92.12 | 60.18 | -91.25 | 61.05 | -92.92 | 59.38 | -92.07 | 60.23 | -93.56 | 58.74 | -83.83 | 68.47 | -79.5 | 72.8 | -81.6 | 70.7
| 59362 | 59371 | BP1J | BP21 | BIPM | 59362-59371 | BP1J-BP21 | -92.1 | 0.1 | -91.05 | 0.1 | -92.91 | 0.1 | -91.9 | 0.1 | -93.5 | 0.1 | -83.71 | 0.1 | -79.65 | 0.5 | -81.62 | 0.2 |  |  |  |  | BP1J-BP21 | 59362-59371 | 195.65 | 43.27 |  | -92.1 | 60.28 | -91.05 | 61.33 | -92.91 | 59.47 | -91.9 | 60.48 | -93.5 | 58.88 | -83.71 | 68.67 | -79.65 | 72.73 | -81.62 | 70.76
Results for Travel system #1 | [29 empty] | Misclosure |  |  |  | 0.1 |  | 0.28 |  | 0.09 |  | 0.25 |  | 0.14 |  | 0.2 |  | 0.07 |  | 0.06
|  |  | BP1J | BP21 | BIPM | [22 empty] | BP1J-BP21 |  | Mean (chosen value) |  |  |  | 60.23 |  | 61.19 |  | 59.425 |  | 60.355 |  | 58.81 |  | 68.57 |  | 72.765 |  | 70.73
[30 empty] | Mean (arithmetic) |  |  |  | 60.23 |  | 61.19 |  | 59.425 |  | 60.355 |  | 58.81 |  | 68.57 |  | 72.765 |  | 70.73
| 59261 | 59266 | BP25 | BP21 | BIPM | 59261-59266 | BP25-BP21 | 11.9 | 0.1 | 17.29 | 0.1 | 10.47 | 0.1 | 15.82 | 0.1 | 10.09 | 0.1 | 16.41 | 0.1 | 13.71 | 0.1 | 19.28 | 0.1 | 18.56 | 0.2 | 19.07 | 0.2 | BP25-BP21 | 59261-59266 | 53.42 | 43.25 |  | 11.9 | 22.07 | 17.29 | 27.46 | 10.47 | 20.64 | 15.82 | 25.99 | 10.09 | 20.26 | 16.41 | 26.58 | 13.71 | 23.88 | 19.28 | 29.45 | 18.56 | 28.73 | 19.07 | 29.24
| 59362 | 59371 | BP25 | BP21 | BIPM | 59362-59371 | BP25-BP21 | 11.95 | 0.1 | 17.47 | 0.1 | 10.56 | 0.1 | 16.02 | 0.1 | 10.15 | 0.1 | 16 | 0.1 | 14.03 | 0.1 | 19.46 | 0.1 | 18.67 | 0.2 | 18.91 | 0.2 | BP25-BP21 | 59362-59371 | 53.44 | 43.27 |  | 11.95 | 22.12 | 17.47 | 27.64 | 10.56 | 20.73 | 16.02 | 26.19 | 10.15 | 20.32 | 16 | 26.17 | 14.03 | 24.2 | 19.46 | 29.63 | 18.67 | 28.84 | 18.91
Results for Travel system #2 | [29 empty] | Misclosure |  |  |  | 0.05 |  | 0.18 |  | 0.09 |  | 0.2 |  | 0.06 |  | 0.41 |  | 0.32 |  | 0.18 |  | 0.11 |  | 0
|  |  | BP25 | BP21 | BIPM | [22 empty] | BP25-BP21 |  | Mean (chosen value) |  |  |  | 22.095 |  | 27.55 |  | 20.685 |  | 26.09 |  | 20.29 |  | 26.375 |  | 24.04 |  | 29.54 |  | 28.785 |  | 29.24
[30 empty] | Mean (arithmetic) |  |  |  | 22.095 |  | 27.55 |  | 20.685 |  | 26.09 |  | 20.29 |  | 26.375 |  | 24.04 |  | 29.54 |  | 28.785 |  | 29.24
BP1K check
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | [14 empty] | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5a
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59261 | 59266 | BP1K | BP21 | BIPM | 59261-59266 | BP1K-BP21 | 9.01 | 0.1 | 9999.99 | 99 | 8.41 | 0.1 | 9.72 | 0.1 | 8.7 | 0.1 | 6.05 | 0.1 | [8 empty] | BP1K-BP21 | 59261-59266 | -33.62 | 43.25 |  | 9.01 | -67.86 | 9999.99 |  | 8.41 | -68.46 | 9.72 | -67.15 | 8.7 | -68.17 | 6.05 | -70.82
| 59362 | 59371 | BP1K | BP21 | BIPM | 59362-59371 | BP1K-BP21 | 8.76 | 0.1 | 10.15 | 0.1 | 8.15 | 0.1 | 9.65 | 0.1 | 8.44 | 0.1 | 5.84 | 0.1 | [8 empty] | BP1K-BP21 | 59362-59371 | -33.62 | 43.26 |  | 8.76 | -68.12 | 10.15 | -66.73 | 8.15 | -68.73 | 9.65 | -67.23 | 8.44 | -68.44 | 5.84 | -71.04
Results for BP1K | [29 empty] | Misclosure |  |  |  | 0.26 |  | 0 |  | 0.27 |  | 0.08 |  | 0.27 |  | 0.22
|  |  | BP1K | BP21 | BIPM | [22 empty] | BP1K-BP21 |  | Mean (chosen value) |  |  |  | -67.99 |  | -66.73 |  | -68.595 |  | -67.19 |  | -68.305 |  | -70.93
[30 empty] | Mean (arithmetic) |  |  |  | -67.99 |  | -66.73 |  | -68.595 |  | -67.19 |  | -68.305 |  | -70.93
BP1C check
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | [14 empty] | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1
| 59261 | 59266 | BP1C | BP21 | BIPM | 59261-59266 | BP1C-BP21 | -182.79 | 0.1 | -176.11 | 0.1 | -183.15 | 0.1 | -177.11 | 0.1 | [12 empty] | BP1C-BP21 | 59261-59266 | 235.54 | 43.25 |  | -182.79 | 9.5 | -176.11 | 16.18 | -183.15 | 9.14 | -177.11 | 15.18
| 59192 | 59198 | BP1C | BP21 | BIPM | 59192-59198 | BP1C-BP21 | -182.82 | 0.1 | -176.12 | 0.1 | -183.27 | 0.1 | -177.23 | 0.1
Results for BP1C
|  |  | BP1C | BP21 | BIPM

## Phase 1
Phase 1 of the calibration trip is BIPM-NICT-TL-BIPM and two traveling systems.
19 June 2020: Opening BIPM; 16 September 2020: NICT; 16 November: Preliminary for TL; 17 December: Closure (preliminary); 25 January 2021: Closure (final); 23 February: extended to GLO C-code and BDS2 for NC5S
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [21 empty] | Reference delays measured in the actual set-up. | [24 empty] | Measured / known antenna cable delays
[30 empty] | The column "Note" is present to match the Table in the report | [24 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed. | [77 empty] | BP21 values are 1001-2020 for GPS, ESA absolute values for GAL, GLO, BDS
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [22 empty] | The colums below make up Table 4 (Traveling - Visited) | [24 empty] | The colums below make up Table 5 (Visited - Reference) | [24 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known)
[78 empty] | BP21 reference values 1001-2020 |  | 28.4 | 27.3 | 30.6 | 0 | 30.7 | 30.9 | 27.7 | 28.6 | 24.4 | 25.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | D2 |  | B1 |  | B2 |  | Pair | Date | CABDLYV | CABDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | D2 |  | B1 |  | B2 |  | Pair | Date | P1 (ns) | P2 (ns) | C1 (ns) | C2(ns) | E1 (ns) | E5 (ns) | D1(ns) | D2(ns) | B1(ns) | B2(ns)
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | INTDLYV | INTDLYV | [6 empty] | Type | (CAB-REF)apriori | TOTDLY(P3) | C1-P1 (result calib) | C1-P1 raw bias | C1-P1 cor bias | Delta(C1-P1)
| 59079 | 59102 | BP1J | NC5S | NICT | 59079-59102 | BP1J-NC5S | -243.09 | 0.1 | -242.39 | 0.1 | -244.02 | 0.1 | -243.43 | 0.1 | -244.64 | 0.1 | -234.47 | 0.1 | -228.38 | 0.5 | -238.19 | 0.5 |  |  |  |  | BP1J-NC5S | 59079-59102 | 403.72 | 266.8 |  | -243.09 | -106.17 | -242.39 | -105.47 | -244.02 | -107.1 | -243.43 | -106.51 | -244.64 | -107.72 | -234.47 | -97.55 | -228.38 | -91.46 | -238.19 | -101.27 |  |  |  |  | NC5S-BP21 via BP1J | 2020.6644 | 0 | 140.8 |  | 224.22 | 365.02 | 224.5 | 365.3 | 224.33 | 365.13 | 224.69 | 365.49 | 224.31 | 365.11 | 223.68 | 364.48 | 221.735 | 362.535 | 229.83 | 370.63 |  |  |  |  | NC5S wrt BP21 via BP1J | 2020.6644 | 393.42 | 392.6 | 395.73 |  | 395.81 | 395.38 | 390.235 | 399.23 |  |  | PolaRx5 (Compensation OFF) | 0 | 127.8869 | 2.31 | 2.79 |  | 2.31
| 59079 | 59102 | BP25 | NC5S | NICT | 59079-59102 | BP25-NC5S | -190.97 | 0.1 | -185.64 | 0.1 | -192.34 | 0.1 | -187.07 | 0.1 | -192.87 | 0.1 | -186.29 | 0.1 | -186.07 | 0.1 | -188.98 | 0.1 | -183.44 | 0.1 | -183.59 | 0.1 | BP25-NC5S | 59079-59102 | 255.61 | 266.8 |  | -190.97 | -202.16 | -185.64 | -196.83 | -192.34 | -203.53 | -187.07 | -198.26 | -192.87 | -204.06 | -186.29 | -197.48 | -186.07 | -197.26 | -188.98 | -200.17 | -183.44 | -194.63 | -183.59 | -194.78 | NC5S-BP21 via BP25 | 2020.6644 | 0 | 140.8 |  | 224.235 | 365.035 | 224.425 | 365.225 | 224.175 | 364.975 | 224.36 | 365.16 | 224.355 | 365.155 | 223.99 | 364.79 | 221.475 | 362.275 | 229.65 | 370.45 | 223.475 | 364.275 | 223.775 | 364.575 | NC5S wrt BP21 via BP25 | 2020.6644 | 393.435 | 392.525 | 395.575 |  | 395.855 | 395.69 | 389.975 | 399.05 | 388.675 | 390.075 | PolaRx5 (Compensation OFF) | 0 | 128.0409 | 2.14 | 2.71 |  | 2.14
| 59079 | 59102 | BP1J | NC5G | NICT | 59079-59102 | BP1J-NC5G | -116.43 | 0.1 | -116.41 | 0.1 | -115.05 | 0.1 | [14 empty] | BP1J-NC5G | 59079-59102 | 403.72 | 171.3 |  | -116.43 | 115.99 | -116.41 | 116.01 | -115.05 | 117.37 | [14 empty] | NC5G-BP21 via BP1J | 2020.6644 | 268.7 | 140.8 | * | 2.06 | -125.84 | 3.02 | -124.88 | -0.14 | -128.04 | [14 empty] | NC5G wrt BP21 via BP1J | 2020.6644 | 1.06 | 0.92 | 1.06 | [7 empty] | GTR50 | 98.5 | 0.1763 | 0 | -7.73 |  | 0
| 59079 | 59102 | BP25 | NC5G | NICT | 59079-59102 | BP25-NC5G | -64.27 | 0.1 | -59.67 | 0.1 | -63.37 | 0.1 | [14 empty] | BP25-NC5G | 59079-59102 | 255.61 | 171.3 |  | -64.27 | 20.04 | -59.67 | 24.64 | -63.37 | 20.94 | [14 empty] | NC5G-BP21 via BP25 | 2020.6644 | 268.7 | 140.8 | * | 2.035 | -125.865 | 2.955 | -124.945 | -0.295 | -128.195 | [14 empty] | NC5G wrt BP21 via BP25 | 2020.6644 | 1.035 | 0.855 | 0.905 | [7 empty] | GTR50 | 98.5 | 0.2131 | -0.13 | -7.8 |  | -0.13
| 59079 | 59102 | BP1J | NC4S | NICT | 59079-59102 | BP1J-NC4S | -78.13 | 0.1 | -77.76 | 0.1 | -78.08 | 0.1 | -77.73 | 0.1 | -78.12 | 0.1 | -77.41 | 0.1 | -77.26 | 0.1 | -78.27 | 0.1 |  |  |  |  | BP1J-NC4S | 59079-59102 | 403.72 | 315.7 |  | -78.13 | 9.89 | -77.76 | 10.26 | -78.08 | 9.94 | -77.73 | 10.29 | -78.12 | 9.9 | -77.41 | 10.61 | -77.26 | 10.76 | -78.27 | 9.75 |  |  |  |  | NC4S-BP21 via BP1J | 2020.6644 | 0 | 140.8 |  | 108.16 | 248.96 | 108.77 | 249.57 | 107.29 | 248.09 | 107.89 | 248.69 | 106.69 | 247.49 | 115.52 | 256.32 | 119.515 | 260.315 | 118.81 | 259.61 |  |  |  |  | NC4S wrt BP21 via BP1J | 2020.6644 | 277.36 | 276.87 | 278.69 |  | 278.19 | 287.22 | 288.015 | 288.21 |  |  | PolaRx4 | 0 | -37.5829 | 1.33 | 1.12 |  | 1.33
| 59079 | 59102 | BP25 | NC4S | NICT | 59079-59102 | BP25-NC4S | -25.99 | 0.1 | -21.02 | 0.1 | -26.4 | 0.1 | -21.37 | 0.1 | -26.34 | 0.1 | -29.25 | 0.1 | -34.78 | 0.5 | -29.33 | 0.5 |  |  |  |  | BP25-NC4S | 59079-59102 | 255.61 | 315.7 |  | -25.99 | -86.08 | -21.02 | -81.11 | -26.4 | -86.49 | -21.37 | -81.46 | -26.34 | -86.43 | -29.25 | -89.34 | -34.78 | -94.87 | -29.33 | -89.42 |  |  |  |  | NC4S-BP21 via BP25 | 2020.6644 | 0 | 140.8 |  | 108.155 | 248.955 | 108.705 | 249.505 | 107.135 | 247.935 | 107.56 | 248.36 | 106.725 | 247.525 | 115.85 | 256.65 | 119.085 | 259.885 | 118.9 | 259.7 |  |  |  |  | NC4S wrt BP21 via BP25 | 2020.6644 | 277.355 | 276.805 | 278.535 |  | 278.225 | 287.55 | 287.585 | 288.3 |  |  | PolaRx4 | 0 | -37.4953 | 1.18 | 1.05 |  | 1.18
* NC5G is a GTR50 with REFDLY=170.2 ns and CABDLY=268.7 ns introduced a priori.
| 59153 | 59165 | BP1J | TLT0 | TL | 59153-59165 | BP1J-TLT0 | 72.17 | 0.1 | 69.78 | 0.1 | 72.22 | 0.1 | 69.99 | 0.1 | 71.9 | 0.1 | 96.31 | 0.5 | [8 empty] | BP1J-TLT0 | 59153-59165 | 192.7 | 157.92 |  | 72.17 | 106.95 | 69.78 | 104.56 | 72.22 | 107 | 69.99 | 104.77 | 71.9 | 106.68 | 96.31 | 131.09 | [8 empty] | TLT0-BP21 via BP1J | 2020.8521 | 119.8 | 140.8 |  | 11.1 | 32.1 | 14.47 | 35.47 | 10.23 | 31.23 | 13.41 | 34.41 | 9.91 | 30.91 | -4.96 | 16.04 | [8 empty] | TLT0 wrt BP21 via BP1J | 2020.8521 | 60.5 | 62.77 | 61.83 |  | 61.61 | 46.94 |  |  |  |  | PolaRx4 | 0 | 18.8729 | 1.33 | 1.37 |  | 1.33
| 59153 | 59165 | BP25 | TLT0 | TL | 59153-59165 | BP25-TLT0 | 115.56 | 0.1 | 117.7 | 0.1 | 115.09 | 0.1 | 117.4 | 0.1 | 114.92 | 0.1 | 135.75 | 0.5 | [8 empty] | BP25-TLT0 | 59153-59165 | 53.51 | 157.92 |  | 115.56 | 11.15 | 117.7 | 13.29 | 115.09 | 10.68 | 117.4 | 12.99 | 114.92 | 10.51 | 135.75 | 31.34 | [8 empty] | TLT0-BP21 via BP25 | 2020.8521 | 119.8 | 140.8 |  | 10.925 | 31.925 | 14.305 | 35.305 | 9.965 | 30.965 | 13.11 | 34.11 | 9.785 | 30.785 | -4.83 | 16.17 | [8 empty] | TLT0 wrt BP21 via BP25 | 2020.8521 | 60.325 | 62.605 | 61.565 |  | 61.485 | 47.07 |  |  |  |  | PolaRx4 | 0 | 18.6824 | 1.24 | 1.39 |  | 1.24
| 59153 | 59165 | BP1J | TLT1 | TL | 59153-59165 | BP1J-TLT1 | -320.53 | 0.1 | -329.48 | 0.1 | -318.88 | 0.1 | [14 empty] | BP1J-TLT1 | 59153-59165 | 192.7 | 0 |  | -320.53 | -127.83 | -329.48 | -136.78 | -318.88 | -126.18 | [14 empty] | TLT1-BP21 via BP1J | 2020.8521 | 0 | 140.8 |  | 245.88 | 386.68 | 255.81 | 396.61 | 243.41 | 384.21 | [14 empty] | TLT1 wrt BP21 via BP1J | 2020.8521 | 415.08 | 423.91 | 414.81 | [7 empty] | Ashtech Z12T | 0 | 401.4376 | -0.27 | -1.69 |  | -0.27
| 59153 | 59165 | BP25 | TLT1 | TL | 59153-59165 | BP25-TLT1 | -277.14 | 0.1 | -281.59 | 0.1 | -276.01 | 0.1 | [14 empty] | BP25-TLT1 | 59153-59165 | 53.51 | 0 |  | -277.14 | -223.63 | -281.59 | -228.08 | -276.01 | -222.5 | [14 empty] | TLT1-BP21 via BP25 | 2020.8521 | 0 | 140.8 |  | 245.705 | 386.505 | 255.675 | 396.475 | 243.145 | 383.945 | [14 empty] | TLT1 wrt BP21 via BP25 | 2020.8521 | 414.905 | 423.775 | 414.545 | [7 empty] | Ashtech Z12T | 0 | 401.2008 | -0.36 | -1.72 |  | -0.36
| 59153 | 59165 | BP1J | TLT3 | TL | 59153-59165 | BP1J-TLT3 | 12.34 | 0.1 | 8.58 | 0.1 | 12.57 | 0.1 | [14 empty] | BP1J-TLT3 | 59153-59165 | 192.7 | 25.64 |  | 12.34 | 179.4 | 8.58 | 175.64 | 12.57 | 179.63 | [14 empty] | TLT3-BP21 via BP1J | 2020.8521 | 143.6 | 140.8 |  | -61.35 | -64.15 | -56.61 | -59.41 | -62.4 | -65.2 | [14 empty] | TLT3 wrt BP21 via BP1J | 2020.8521 | -35.75 | -32.11 | -34.6 | [7 empty] | GTR50 | 0 | 76.5862 | 1.15 | 1.19 |  | 1.15
| 59153 | 59165 | BP25 | TLT3 | TL | 59153-59165 | BP25-TLT3 | 55.72 | 0.1 | 56.44 | 0.1 | 55.44 | 0.1 | [14 empty] | BP25-TLT3 | 59153-59165 | 53.51 | 25.64 |  | 55.72 | 83.59 | 56.44 | 84.31 | 55.44 | 83.31 | [14 empty] | TLT3-BP21 via BP25 | 2020.8521 | 143.6 | 140.8 |  | -61.515 | -64.315 | -56.715 | -59.515 | -62.665 | -65.465 | [14 empty] | TLT3 wrt BP21 via BP25 | 2020.8521 | -35.915 | -32.215 | -34.865 | [7 empty] | GTR50 | 0 | 76.3285 | 1.05 | 1.18 |  | 1.05
| 59153 | 59165 | BP1J | TLT5 | TL | 59153-59165 | BP1J-TLT5 | -109.5 | 0.1 | -108.54 | 0.1 | -110.36 | 0.1 |  |  | -110.96 | 0.1 | -98.83 | 0.1 | [8 empty] | BP1J-TLT5 | 59153-59165 | 192.7 | 0 | * | -109.5 | 83.2 | -108.54 | 84.16 | -110.36 | 82.34 |  |  | -110.96 | 81.74 | -98.83 | 93.87 | [8 empty] | TLT5-BP21 via BP1J | 2020.8521 | 0 | 140.8 | * | 34.85 | 175.65 | 34.87 | 175.67 | 34.89 | 175.69 |  |  | 34.85 | 175.65 | 32.26 | 173.06 | [8 empty] | TLT5 wrt BP21 via BP1J | 2020.8521 | 204.05 | 202.97 | 206.29 |  | 206.35 | 203.96 |  |  |  |  | PolaRx5 | 0 | 205.7186 | 2.24 | -1.69 |  | 2.24
| 59153 | 59165 | BP25 | TLT5 | TL | 59153-59165 | BP25-TLT5 | -66.13 | 0.1 | -60.64 | 0.1 | -67.46 | 0.1 |  |  | -67.94 | 0.1 | -59.61 | 0.1 | [8 empty] | BP25-TLT5 | 59153-59165 | 53.51 | 0 | * | -66.13 | -12.62 | -60.64 | -7.13 | -67.46 | -13.95 |  |  | -67.94 | -14.43 | -59.61 | -6.1 | [8 empty] | TLT5-BP21 via BP25 | 2020.8521 | 0 | 140.8 | * | 34.695 | 175.495 | 34.725 | 175.525 | 34.595 | 175.395 |  |  | 34.725 | 175.525 | 32.61 | 173.41 | [8 empty] | TLT5 wrt BP21 via BP25 | 2020.8521 | 203.895 | 202.825 | 205.995 |  | 206.225 | 204.31 |  |  |  |  | PolaRx5 | 0 | 205.5481 | 2.1 | -1.72 |  | 2.1
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [22 empty] | The colums below make up Table 3 (Traveling - Reference)
BP21 reference (Auto-compensation OFF 59010-59015, ON after 59023: ON to be used)
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | D2 |  | B1 |  | B2
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59010 | 59015 | BP1J | BP21 | BIPM | 59010-59015 | BP1J-BP21 | -4.64 | 0.1 | -3.67 | 0.1 | -5.46 | 0.1 | -4.52 | 0.1 | -6.08 | 0.1 | 3.33 | 0.1 | [8 empty] | BP1J-BP21 | 59010-59015 | 194.64 | 72.82 |  | -4.64 | 117.18 | -3.67 | 118.15 | -5.46 | 116.36 | -4.52 | 117.3 | -6.08 | 115.74 | 3.33 | 125.15
| 59024 | 59029 | BP1J | BP21 | BIPM | 59024-59029 | BP1J-BP21 | -34.29 | 0.1 | -33.36 | 0.1 | -35.12 | 0.1 | -34.23 | 0.1 | -35.76 | 0.1 | -26.1 | 0.1 | -22.09 | 0.5 | -23.73 | 0.5 |  |  |  |  | BP1J-BP21 | 59024-59029 | 195.69 | 43.25 |  | -34.29 | 118.15 | -33.36 | 119.08 | -35.12 | 117.32 | -34.23 | 118.21 | -35.76 | 116.68 | -26.1 | 126.34 | -22.09 | 130.35 | -23.73 | 128.71
| 59192 | 59198 | BP1J | BP21 | BIPM | 59192-59198 | BP1J-BP21 | -34.48 | 0.1 | -33.45 | 0.1 | -35.29 | 0.1 | -34.28 | 0.1 | -35.93 | 0.1 | -26.51 | 0.1 | -22.23 | 0.2 | -24.02 | 0.2 |  |  |  |  | BP1J-BP21 | 59192-59198 | 195.69 | 43.26 |  | -34.48 | 117.95 | -33.45 | 118.98 | -35.29 | 117.14 | -34.28 | 118.15 | -35.93 | 116.5 | -26.51 | 125.92 | -22.23 | 130.2 | -24.02 | 128.41
Results for Travel system #1 | [29 empty] | Misclosure |  |  |  | 0.2 |  | 0.1 |  | 0.18 |  | 0.06 |  | 0.18 |  | 0.42 |  | 0.15 |  | 0.3
|  |  | BP1J | BP21 | BIPM | [22 empty] | BP1J-BP21 |  | Mean (chosen value) |  |  |  | 118.05 |  | 119.03 |  | 117.23 |  | 118.18 |  | 116.59 |  | 126.13 |  | 130.275 |  | 128.56
[30 empty] | Mean (arithmetic) |  |  |  | 118.05 |  | 119.03 |  | 117.23 |  | 118.18 |  | 116.59 |  | 126.13 |  | 130.275 |  | 128.56
| 59010 | 59015 | BP25 | BP21 | BIPM | 59010-59015 | BP25-BP21 | 41.55 | 0.1 | 47.06 | 0.1 | 40.1 | 0.1 | 45.59 | 0.1 | 39.78 | 0.1 | 45.75 | 0.1 | [8 empty] | BP25-BP21 | 59010-59015 | 52.6 | 72.82 |  | 41.55 | 21.33 | 47.06 | 26.84 | 40.1 | 19.88 | 45.59 | 25.37 | 39.78 | 19.56 | 45.75 | 25.53
| 59024 | 59029 | BP25 | BP21 | BIPM | 59024-59029 | BP25-BP21 | 11.86 | 0.1 | 17.55 | 0.1 | 10.4 | 0.1 | 16.07 | 0.1 | 10.08 | 0.1 | 16.15 | 0.1 | 14.36 | 0.2 | 19.27 | 0.2 | 18.67 | 0.2 | 18.77 | 0.2 | BP25-BP21 | 59024-59029 | 53.51 | 43.25 |  | 11.86 | 22.12 | 17.55 | 27.81 | 10.4 | 20.66 | 16.07 | 26.33 | 10.08 | 20.34 | 16.15 | 26.41 | 14.36 | 24.62 | 19.27 | 29.53 | 18.67 | 28.93 | 18.77 | 29.03
| 59192 | 59198 | BP25 | BP21 | BIPM | 59192-59198 | BP25-BP21 | 11.89 | 0.1 | 17.24 | 0.1 | 10.49 | 0.1 | 15.73 | 0.1 | 10.11 | 0.1 | 16.47 | 0.1 | 13.67 | 0.1 | 19.29 | 0.1 | 18.62 | 0.1 | 18.82 | 0.1 | BP25-BP21 | 59192-59198 | 53.4 | 43.26 |  | 11.89 | 22.03 | 17.24 | 27.38 | 10.49 | 20.63 | 15.73 | 25.87 | 10.11 | 20.25 | 16.47 | 26.61 | 13.67 | 23.81 | 19.29 | 29.43 | 18.62 | 28.76 | 18.82 | 28.96
Results for Travel system #2 | [29 empty] | Misclosure |  |  |  | 0.09 |  | 0.43 |  | 0.03 |  | 0.46 |  | 0.09 |  | 0.2 |  | 0.81 |  | 0.1 |  | 0.17 |  | 0.07
|  |  | BP25 | BP21 | BIPM | [22 empty] | BP25-BP21 |  | Mean (chosen value) |  |  |  | 22.075 |  | 27.595 |  | 20.645 |  | 26.1 |  | 20.295 |  | 26.51 |  | 24.215 |  | 29.48 |  | 28.845 |  | 28.995
[30 empty] | Mean (arithmetic) |  |  |  | 22.075 |  | 27.595 |  | 20.645 |  | 26.1 |  | 20.295 |  | 26.51 |  | 24.215 |  | 29.48 |  | 28.845 |  | 28.995
BP1C check
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59010 | 59015 | BP1C | BP21 | BIPM | 59010-59015 | BP1C-BP21 | -153.17 | 0.1 | -146.5 | 0.1 | -153.6 | 0.1 | -147.57 | 0.1 | [12 empty] | BP1C-BP21 | 59010-59015 | 234.99 | 72.82 |  | -153.17 | 9 | -146.5 | 15.67 | -153.6 | 8.57 | -147.57 | 14.6
| 59192 | 59198 | BP1C | BP21 | BIPM | 59192-59198 | BP1C-BP21 | -182.79 | 0.1 | -176.16 | 0.1 | -183.16 | 0.1 | -177.16 | [13 empty] | BP1C-BP21 | 59192-59198 | 235.71 | 43.26 |  | -182.79 | 9.66 | -176.16 | 16.29 | -183.16 | 9.29 | -177.16 | 15.29
Results for BP1C | [29 empty] | Misclosure |  |  |  | 0.66 |  | 0.62 |  | 0.72 |  | 0.69
|  |  | BP1C | BP1J | BIPM | [22 empty] | BP1C-BP1J |  | Mean (chosen value) |  |  |  | 9.33 |  | 15.98 |  | 8.93 |  | 14.945
[30 empty] | Mean (arithmetic) |  |  |  | 9.33 |  | 15.98 |  | 8.93 |  | 14.945
BP1K check
[8 empty] | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(D2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc
| 59010 | 59015 | BP1K | BP21 | BIPM | 59010-59015 | BP1K-BP21 | 38.2 | 0.1 | 39.36 | 0.1 | 37.59 | 0.1 | 38.74 | 0.1 | 37.83 | 0.1 | 35.26 | 0.1 | [8 empty] | BP1K-BP21 | 59010-59015 | -32.9 | 72.82 |  | 38.2 | -67.52 | 39.36 | -66.36 | 37.59 | -68.13 | 38.74 | -66.98 | 37.83 | -67.89 | 35.26 | -70.46
| 59024 | 59029 | BP1K | BP21 | BIPM | 59024-59029 | BP1K-BP21 | 8.48 | 0.1 | 9.63 | 0.1 | 7.85 | 0.1 | 9 | 0.1 | 8.06 | 0.1 | 5.53 | 0.1 | [8 empty] | BP1K-BP21 | 59024-59029 | -32.15 | 43.25 |  | 8.48 | -66.92 | 9.63 | -65.77 | 7.85 | -67.55 | 9 | -66.4 | 8.06 | -67.34 | 5.53 | -69.87
| 59192 | 59198 | BP1K | BP21 | BIPM | 59192-59198 | BP1K-BP21 | 8.53 | 0.1 | 9.76 | 0.1 | 7.91 | 0.1 | 9.18 | 0.1 | 8.15 | 0.1 | 5.61 | 0.1 | [8 empty] | BP1K-BP21 | 59192-59198 | -32.15 | 43.26 |  | 8.53 | -66.88 | 9.76 | -65.65 | 7.91 | -67.5 | 9.18 | -66.23 | 8.15 | -67.26 | 5.61 | -69.8
Results for BP1K | [29 empty] | Misclosure |  |  |  | 0.04 |  | 0.12 |  | 0.05 |  | 0.17 |  | 0.08 |  | 0.07
|  |  | BP1K | BP21 | BIPM | [22 empty] | BP1K-BP21 |  | Mean (chosen value) |  |  |  | -66.9 |  | -65.71 |  | -67.525 |  | -66.315 |  | -67.3 |  | -69.835
[30 empty] | Mean (arithmetic) |  |  |  | -66.9 |  | -65.71 |  | -67.525 |  | -66.315 |  | -67.3 |  | -69.835

## Phase 1B
Phase 1B of the calibration trip is BIPM-NIM-BIPM and two traveling systems.
29 June 2020: Opening BIPM; 27 October: NIM results; 17 December: Closure (preliminary); 25 January 2021: Closure (final)
Sections with the following colors are to be filled by the user. The data will be used to generate the tables for the report.
| Generic information | [6 empty] | Raw code differences | [21 empty] | Reference delays measured in the actual set-up. | [24 empty] | Measured / known antenna cable delays
[30 empty] | The column "Note" is present to match the Table in the report | [24 empty] | The column "Note" is present to match the Table in the report
First part is for the legs, excluding the laboratory where the closure takes place: Use the area between red lines, one line for each measurement. Lines can be added as needed. | [77 empty] | BP21 values are 1001-2020 for GPS, ESA absolute values for GAL, GLO, BDS
[5 empty] | The colums below make up Table 2.2: Raw code differences (Traveling - Visited) | [22 empty] | The colums below make up Table 4 (Traveling - Visited) | [24 empty] | The colums below make up Table 5 (Visited - Reference) | [24 empty] | The colums below make up Table 6 (To be used if Reference INTDLY values are known) | [8 empty] | 28.6 D2
[78 empty] | BP21 reference values 1001-2018 |  | 28.4 | 27.3 | 30.6 |  | 30.7 | 30.9 | 27.7 | 27.5 | 24.4 | 25.5
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(R2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYV | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | R2 |  | B1 |  | B2 |  | Pair | Date | CABDLYV | CABDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | R2 |  | B1 |  | B2 |  | Pair | Date | P1 (ns) | P2 (ns) | C1 (ns) | C2 | E1 (ns) | E5a (ns) | D1(ns) | R2(ns) | B1(ns) | B2(ns)
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [5 empty] | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY | DSYSDLY | DINTDLY |  |  | INTDLYV | INTDLYV | INTDLYV |  | INTDLYV | INTDLYV | INTDLYV | INTDLYV | INTDLYV | INTDLYV | Type | (CAB-REF)apriori | TOTDLY(P3) | C1-P1 (result calib) | C1-P1 raw bias | C1-P1 cor bias | Delta(C1-P1)
| 59100 | 59106 | TS03 | IM15 | NIM | 59100-59106 | TS03-IM15 | -117.2 | 0.1 | -109.42 | 0.1 | -116.14 | 0.1 |  |  | -115.53 | 0.1 | -103.24 | 0.1 | -116.83 | 0.5 | -103.31 | 0.2 | -122.9 | 0.1 | -119.42 |  | TS03-IM15 | 59100-59106 | 283.2 | 119.4 |  | -117.2 | 46.6 | -109.42 | 54.38 | -116.14 | 47.66 |  |  | -115.53 | 48.27 | -103.24 | 60.56 | -116.83 | 46.97 | -103.31 | 60.49 | -122.9 | 40.9 | -119.42 | 44.38 | IM15-BP21 via TS03 | 2020.6986 | 212.4 | 140.8 |  | 15.405 | -56.195 | 6.085 | -65.515 | 14.31 | -57.29 |  |  | 13.685 | -57.915 | 2.75 | -68.85 | 15.675 | -55.925 | 2.27 | -69.33 | 20.5 | -51.1 | 16.455 | -55.145 | IM15 wrt BP21 via TS03 | 2020.726 | -27.82 | -38.235 | -26.745 |  | -27.265 | -36.61 | -28.225 | -41.83 | -26.7 | -29.645 | NIM-TF | 0 | 81.2712 | 1.075 | 1.37 |  | 1.075
| 59100 | 59106 | TS04 | IM15 | NIM | 59100-59106 | TS04-IM15 | -114.91 | 0.1 | -107.42 | 0.1 | -113.84 | 0.1 |  |  | -113.22 | 0.1 | -101.04 | 0.1 | -113.84 | 0.5 | -100.86 | 0.2 | -121.55 | 0.1 | -117.81 |  | TS04-IM15 | 59100-59106 | 284 | 119.4 |  | -114.91 | 49.69 | -107.42 | 57.18 | -113.84 | 50.76 |  |  | -113.22 | 51.38 | -101.04 | 63.56 | -113.84 | 50.76 | -100.86 | 63.74 | -121.55 | 43.05 | -117.81 | 46.79 | IM15-BP21 via TS04 | 2020.6986 | 212.4 | 140.8 |  | 15.855 | -55.745 | 5.975 | -65.625 | 14.69 | -56.91 |  |  | 14.01 | -57.59 | 3.105 | -68.495 | 15.535 | -56.065 | 2.635 | -68.965 | 21.235 | -50.365 | 16.325 | -55.275 | IM15 wrt BP21 via TS04 | 2020.726 | -27.315 | -38.33 | -26.33 |  | -26.9 | -36.255 | -28.365 | -41.465 | -25.965 | -29.775 | NIM-TF | 0 | 82.7032 | 0.985 | 1.39 |  | 0.985
| 59100 | 59106 | TS03 | IMEJ | NIM | 59100-59106 | TS03-IMEJ | -51.91 | 0.1 | -54.8 | 0.1 | -49.88 | 0.1 | [14 empty] | TS03-IMEJ | 59100-59106 | 283.2 | 121.7 | * | -51.91 | 109.59 | -54.8 | 106.7 | -49.88 | 111.62 | [14 empty] | IMEJ-BP21 via TS03 | 2020.6986 | 248.7 | 140.8 | * | -47.585 | -155.485 | -46.235 | -154.135 | -49.65 | -157.55 | [14 empty] | IMEJ wrt BP21 via TS03 | 2020.726 | -0.16 | 0.135 | -0.02 | [7 empty] | GTR50 | 127 | -0.61578 | 0.14 | -1.69 |  | 0.14
| 59100 | 59106 | TS04 | IMEJ | NIM | 59100-59106 | TS04-IMEJ | -49.63 | 0.1 | -52.77 | 0.1 | -47.58 | 0.1 | [14 empty] | TS04-IMEJ | 59100-59106 | 284 | 121.7 | * | -49.63 | 112.67 | -52.77 | 109.53 | -47.58 | 114.72 | [14 empty] | IMEJ-BP21 via TS04 | 2020.6986 | 248.7 | 140.8 | * | -47.125 | -155.025 | -46.375 | -154.275 | -49.27 | -157.17 | [14 empty] | IMEJ wrt BP21 via TS04 | 2020.726 | 0.36 | 0.02 | 0.405 | [7 empty] | GTR50 | 127 | 0.8853 | 0.045 | -1.72 |  | 0.045
| 59100 | 59106 | TS03 | BJNM | NIM | 59100-59106 | TS03-BJNM | 78.65 | 0.1 | 68.6 | 0.1 | 79.18 | 0.1 |  |  | 78.78 | 0.1 | [10 empty] | TS03-BJNM | 59100-59106 | 283.2 | 324.8 |  | 78.65 | 37.05 | 68.6 | 27 | 79.18 | 37.58 |  |  | 78.78 | 37.18 | [10 empty] | BJNM-BP21 via TS03 | 2020.6986 | 125 | 140.8 |  | 24.955 | 40.755 | 33.465 | 49.265 | 24.39 | 40.19 |  |  | 24.775 | 40.575 | [10 empty] | BJNM wrt BP21 via TS03 | 2020.726 | 69.21 | 76.68 | 70.845 |  | 71.315 | [5 empty] | PolaRx3 | 0 | -142.1312 | 1.635 | 1.37 |  | 1.635
| 59100 | 59106 | TS04 | BJNM | NIM | 59100-59106 | TS04-BJNM | 80.93 | 0.1 | 70.6 | 0.1 | 81.48 | 0.1 |  |  | 81.11 | 0.1 | [10 empty] | TS04-BJNM | 59100-59106 | 284 | 324.8 |  | 80.93 | 40.13 | 70.6 | 29.8 | 81.48 | 40.68 |  |  | 81.11 | 40.31 | [10 empty] | BJNM-BP21 via TS04 | 2020.6986 | 125 | 140.8 |  | 25.415 | 41.215 | 33.355 | 49.155 | 24.77 | 40.57 |  |  | 25.08 | 40.88 | [10 empty] | BJNM wrt BP21 via TS04 | 2020.726 | 69.72 | 76.59 | 71.275 |  | 71.685 | [5 empty] | PolaRx3 | 0 | -140.6941 | 1.555 | 1.39 |  | 1.555
| 59100 | 59106 | TS03 | IM21 | NIM | 59100-59106 | TS03-IM21 | -105.28 | 0.1 | -104.22 | 0.1 | -104.25 | 0.1 |  |  | -103.57 | 0.1 | -98.31 | 0.1 | [5 empty] | 0.1 | -119.11 |  | TS03-IM21 | 59100-59106 | 283.2 | 119.8 |  | -105.28 | 58.12 | -104.22 | 59.18 | -104.25 | 59.15 |  |  | -103.57 | 59.83 | -98.31 | 65.09 | [6 empty] | -119.11 | 44.29 | IM21-BP21 via TS03 | 2020.6986 | 212.3 | 140.8 |  | 3.885 | -67.615 | 1.285 | -70.215 | 2.82 | -68.68 |  |  | 2.125 | -69.375 | -1.78 | -73.28 | [8 empty] | IM21 wrt BP21 via TS03 | 2020.726 | -39.2 | -42.925 | -38.12 |  | -38.71 | -40.955 |  |  |  |  | NIM-TF | 0 | 59.0551 | 1.08 | 1.37 |  | 1.08
| 59100 | 59106 | TS04 | IM21 | NIM | 59100-59106 | TS04-IM21 | -103 | 0.1 | -102.2 | 0.1 | -101.95 | 0.1 |  |  | -101.26 | 0.1 | -96.09 | 0.1 | [5 empty] | 0.1 | -117.02 |  | TS04-IM21 | 59100-59106 | 284 | 119.8 |  | -103 | 61.2 | -102.2 | 62 | -101.95 | 62.25 |  |  | -101.26 | 62.94 | -96.09 | 68.11 | [6 empty] | -117.02 | 47.18 | IM21-BP21 via TS04 | 2020.6986 | 212.3 | 140.8 |  | 4.345 | -67.155 | 1.155 | -70.345 | 3.2 | -68.3 |  |  | 2.45 | -69.05 | -1.445 | -72.945 | [8 empty] | IM21 wrt BP21 via TS04 | 2020.726 | -38.69 | -43.045 | -37.695 |  | -38.345 | -40.625 |  |  |  |  | NIM-TF | 0 | 60.5385 | 0.995 | 1.39 |  | 0.995
| 59100 | 59106 | TS03 | IMEU | NIM | 59100-59106 | TS03-IMEU | -157.37 | 0.1 | -173.08 | 0.1 | -156.55 | 0.1 | [11 empty] | 0.1 |  |  | TS03-IMEU | 59100-59106 | 283.2 | 120.1 |  | -157.37 | 5.73 | -173.08 | -9.98 | -156.55 | 6.55 | [14 empty] | IMEU-BP21 via TS03 | 2020.6986 | 250.3 | 140.8 |  | 56.275 | -53.225 | 70.445 | -39.055 | 55.42 | -54.08 | [14 empty] | IMEU wrt BP21 via TS03 | 2020.726 | -24.805 | -11.715 | -23.48 | [7 empty] | NIM-TF | 0 | 85.171 | 1.325 | 1.37 |  | 1.325
| 59100 | 59106 | TS04 | IMEU | NIM | 59100-59106 | TS04-IMEU | -155.07 | 0.1 | -171.08 | 0.1 | -154.24 | 0.1 | [11 empty] | 0.1 |  |  | TS04-IMEU | 59100-59106 | 284 | 120.1 |  | -155.07 | 8.83 | -171.08 | -7.18 | -154.24 | 9.66 | [14 empty] | IMEU-BP21 via TS04 | 2020.6986 | 250.3 | 140.8 |  | 56.715 | -52.785 | 70.335 | -39.165 | 55.79 | -53.71 | [14 empty] | IMEU wrt BP21 via TS04 | 2020.726 | -24.31 | -11.805 | -23.06 | [7 empty] | NIM-TF | 0 | 86.5698 | 1.25 | 1.39 |  | 1.25
| 59120 | 59126 | TS03 | IM15 | NIM | 59120-59126 | TS03-IM15 | -117.15 | 0.1 | -109.38 | 0.1 | -116.03 | 0.1 |  |  | -115.43 | 0.1 | -103.18 | 0.1 | -116.74 | 0.5 | -103.26 | 0.2 | -123.03 | 0.1 | -119.57 | 0.2 | TS03-IM15 | 59120-59126 | 283.2 | 119.4 |  | -117.15 | 46.65 | -109.38 | 54.42 | -116.03 | 47.77 |  |  | -115.43 | 48.37 | -103.18 | 60.62 |  |  |  |  | -123.03 | 40.77 | -119.57 | 44.23 | IM15-BP21 via TS03 | 2020.7534 | 212.4 | 140.8 |  | 15.355 | -56.245 | 6.045 | -65.555 | 14.2 | -57.4 |  |  | 13.585 | -58.015 | 2.69 | -68.91 | [10 empty] | 0.05 | 0.04 | 0.11 |  | 0.1 | 0.06
| 59120 | 59126 | TS04 | IM15 | NIM | 59120-59126 | TS04-IM15 | -114.97 | 0.1 | -107.41 | 0.1 | -113.8 | 0.1 |  |  | -113.2 | 0.1 | -100.94 | 0.1 | -113.68 | 0.5 | -100.82 | 0.2 | -121.52 | 0.1 | -117.81 | 0.2 | TS04-IM15 | 59120-59126 | 284 | 119.4 |  | -114.97 | 49.63 | -107.41 | 57.19 | -113.8 | 50.8 |  |  | -113.2 | 51.4 | -100.94 | 63.66 |  |  |  |  | -121.52 | 43.08 | -117.81 | 46.79 | IM15-BP21 via TS04 | 2020.7534 | 212.4 | 140.8 |  | 15.915 | -55.685 | 5.965 | -65.635 | 14.65 | -56.95 |  |  | 13.99 | -57.61 | 3.005 | -68.595 | [10 empty] | -0.06 | 0.01 | 0.04 |  | 0.02 | 0.1
| 59120 | 59126 | TS03 | IMEJ | NIM | 59120-59126 | TS03-IMEJ | -51.76 | 0.1 | -54.74 | 0.1 | -49.74 | 0.1 | [14 empty] | TS03-IMEJ | 59120-59126 | 283.2 | 121.7 | * | -51.76 | 109.74 | -54.74 | 106.76 | -49.74 | 111.76 | [14 empty] | IMEJ-BP21 via TS03 | 2020.7534 | 248.7 | 140.8 | * | -47.735 | -155.635 | -46.295 | -154.195 | -49.79 | -157.69 | [16 empty] | 0.15 | 0.06 | 0.14
| 59120 | 59126 | TS04 | IMEJ | NIM | 59120-59126 | TS04-IMEJ | -49.6 | 0.1 | -52.76 | 0.1 | -47.53 | 0.1 | [14 empty] | TS04-IMEJ | 59120-59126 | 284 | 121.7 | * | -49.6 | 112.7 | -52.76 | 109.54 | -47.53 | 114.77 | [14 empty] | IMEJ-BP21 via TS04 | 2020.7534 | 248.7 | 140.8 | * | -47.155 | -155.055 | -46.385 | -154.285 | -49.32 | -157.22 | [16 empty] | 0.03 | 0.01 | 0.05
| 59120 | 59126 | TS03 | BJNM | NIM | 59120-59126 | TS03-BJNM | 78.54 | 0.1 | 68.37 | 0.1 | 79.07 | 0.1 |  |  | 78.7 | 0.1 | [10 empty] | TS03-BJNM | 59120-59126 | 283.2 | 324.8 |  | 78.54 | 36.94 | 68.37 | 26.77 | 79.07 | 37.47 |  |  | 78.7 | 37.1 | [10 empty] | BJNM-BP21 via TS03 | 2020.7534 | 125 | 140.8 |  | 25.065 | 40.865 | 33.695 | 49.495 | 24.5 | 40.3 |  |  | 24.855 | 40.655 | [12 empty] | -0.11 | -0.23 | -0.11 |  | -0.08
| 59120 | 59126 | TS04 | BJNM | NIM | 59120-59126 | TS04-BJNM | 80.72 | 0.1 | 70.33 | 0.1 | 81.27 | 0.1 |  |  | 80.9 | 0.1 | [10 empty] | TS04-BJNM | 59120-59126 | 284 | 324.8 |  | 80.72 | 39.92 | 70.33 | 29.53 | 81.27 | 40.47 |  |  | 80.9 | 40.1 | [10 empty] | BJNM-BP21 via TS04 | 2020.7534 | 125 | 140.8 |  | 25.625 | 41.425 | 33.625 | 49.425 | 24.98 | 40.78 |  |  | 25.29 | 41.09 | [12 empty] | -0.21 | -0.27 | -0.21 |  | -0.21
| 59120 | 59126 | TS03 | IM21 | NIM | 59120-59126 | TS03-IM21 | -105.31 | 0.1 | -104.2 | 0.1 | -104.17 | 0.1 |  |  | -103.5 | 0.1 | -98.42 | 0.1 | [5 empty] | 0.1 | -117.34 | 5 | TS03-IM21 | 59120-59126 | 283.2 | 119.8 |  | -105.31 | 58.09 | -104.2 | 59.2 | -104.17 | 59.23 |  |  | -103.5 | 59.9 | -98.42 | 64.98 | [6 empty] | -117.34 | 46.06 | IM21-BP21 via TS03 | 2020.7534 | 212.3 | 140.8 |  | 3.915 | -67.585 | 1.265 | -70.235 | 2.74 | -68.76 |  |  | 2.055 | -69.445 | -1.67 | -73.17 | [10 empty] | -0.03 | 0.02 | 0.08 |  | 0.07 | -0.11
| 59120 | 59126 | TS04 | IM21 | NIM | 59120-59126 | TS04-IM21 | -103.13 | 0.1 | -102.2 | 0.1 | -101.96 | 0.1 |  |  | -101.27 | 0.1 | -96.15 | 0.1 | [5 empty] | 0.1 | -115.46 | 5 | TS04-IM21 | 59120-59126 | 284 | 119.8 |  | -103.13 | 61.07 | -102.2 | 62 | -101.96 | 62.24 |  |  | -101.27 | 62.93 | -96.15 | 68.05 | [6 empty] | -115.46 | 48.74 | IM21-BP21 via TS04 | 2020.7534 | 212.3 | 140.8 |  | 4.475 | -67.025 | 1.155 | -70.345 | 3.21 | -68.29 |  |  | 2.46 | -69.04 | -1.385 | -72.885 | [10 empty] | -0.13 | 0 | -0.01 |  | -0.01 | -0.06
| 59120 | 59126 | TS03 | IMEU | NIM | 59120-59126 | TS03-IMEU | -157.41 | 0.1 | -173.16 | 0.1 | -156.55 | 0.1 | [11 empty] | 0.1 |  |  | TS03-IMEU | 59120-59126 | 283.2 | 120.1 |  | -157.41 | 5.69 | -173.16 | -10.06 | -156.55 | 6.55 | [14 empty] | IMEU-BP21 via TS03 | 2020.7534 | 250.3 | 140.8 |  | 56.315 | -53.185 | 70.525 | -38.975 | 55.42 | -54.08 | [16 empty] | -0.04 | -0.08 | 0
| 59120 | 59126 | TS04 | IMEU | NIM | 59120-59126 | TS04-IMEU | -155.22 | 0.1 | -171.2 | 0.1 | -154.34 | 0.1 | [11 empty] | 0.1 |  |  | TS04-IMEU | 59120-59126 | 284 | 120.1 |  | -155.22 | 8.68 | -171.2 | -7.3 | -154.34 | 9.56 | [14 empty] | IMEU-BP21 via TS04 | 2020.7534 | 250.3 | 140.8 |  | 56.865 | -52.635 | 70.455 | -39.045 | 55.89 | -53.61 | [16 empty] | -0.15 | -0.12 | -0.1
Second part is for the laboratory where the closure takes place: Use the area between red lines for each traveling system, one line for each measurement. Lines can be added as needed, Misclosure and Mean values are computed for data between two red lines.
[5 empty] | The colums below make up Table 2.1: Raw code differences (Traveling - Reference) | [22 empty] | The colums below make up Table 3 (Traveling - Reference)
BP21 reference. BP21 was in auto-compensation OFF on 59010-59015, ON after 59023. The REFDLY value for 59010-59015 was computed withthe Xo value from the receiver to match auto-compensation ON. "Initial" line refers to early run of dclrinex with inaccurate baseline.
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc | RAWDIF(D1) | Unc | RAWDIF(R2) | Unc | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1 |  | C2 |  | E1 |  | E5 |  | D1 |  | R2 |  | B1 |  | B2
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59010 | 59015 | TS03 | BP21 | BIPM | 59010-59015 | TS03-BP21 | 107.82 | 0.1 | 106.24 | 0.1 | 107.79 | 0.1 | 106.27 | 0.1 | 107.8 | 0.1 | 109 | 0.1 | 108.27 | 0.1 | 108.64 | 0.1 | 107.06 | 0.1 | 106.83 | 0.1 | TS03-BP21 | 59010-59015 | 27.23 | 72.76 | * | 107.82 | 62.29 | 106.24 | 60.71 | 107.79 | 62.26 | 106.27 | 60.74 | 107.8 | 62.27 | 109 | 63.47 | 108.27 | 62.74 | 108.64 | 63.11 | 107.06 | 61.53 | 106.83 | 61.3
| 59191 | 59197 | TS03 | BP21 | BIPM | 59191-59197 | TS03-BP21 | 78.03 | 0.1 | 76.53 | 0.1 | 77.99 | 0.1 | 76.6 |  | 77.95 | 0.1 | 79.46 | 0.1 | 78.86 | 0.1 | 78.72 | 0.1 | 77.58 | 0.1 | 76.68 | 0.1 | TS03-BP21 | 59191-59197 | 26.95 | 43.26 |  | 78.03 | 61.72 | 76.53 | 60.22 | 77.99 | 61.68 | 76.6 | 60.29 | 77.95 | 61.64 | 79.46 | 63.15 | 78.86 | 62.55 | 78.72 | 62.41 | 77.58 | 61.27 | 76.68 | 60.37
Initial run | 59191 | 59197 | TS03 | BP21 | BIPM | 59191-59197 | TS03-BP21 | 77.56 | 0.1 | 76.07 | 0.1 | 77.55 | 0.1 | 76.13 |  | 77.48 | 0.1 | 79.08 | 0.1 |  |  |  |  | 77.05 | 0.1 | 75.89 | 0.1 | * BP21 was in auto-compensation OFF on 59010-59015, ON after 59023. The REFDLY value for 59010-59015 was computed withthe Xo value from the receiver during 59024-59029 to match auto-compensation ON.
"Initial" line refers to early run of dclrinex with inaccurate baseline.
Results for Travel system #1 | [29 empty] | Misclosure |  |  |  | 0.57 |  | 0.49 |  | 0.58 |  | 0.45 |  | 0.63 |  | 0.32 |  | 0.19 |  | 0.7 |  | 0.26 |  | 0.93
|  |  | TS03 | BP21 | BIPM | [22 empty] | TS03-BP21 |  | Mean (chosen value) |  |  |  | 62.005 |  | 60.465 |  | 61.97 |  | 60.515 |  | 61.955 |  | 63.31 |  | 62.645 |  | 62.76 |  | 61.4 |  | 60.835
[30 empty] | Mean (arithmetic) |  |  |  | 62.005 |  | 60.465 |  | 61.97 |  | 60.515 |  | 61.955 |  | 63.31 |  | 62.645 |  | 62.76 |  | 61.4 |  | 60.835
| 59010 | 59015 | TS04 | BP21 | BIPM | 59010-59015 | TS04-BP21 | 110.43 | 0.1 | 108.09 | 0.1 | 110.34 | 0.1 | 108.11 | 0.1 | 110.27 | 0.1 | 111.51 | 0.1 | 111.08 | 0.1 | 111.2 | 0.1 | 109.07 | 0.1 | 107.99 | 0.1 | TS04-BP21 | 59010-59015 | 28.09 | 72.76 | * | 110.43 | 65.76 | 108.09 | 63.42 | 110.34 | 65.67 | 108.11 | 63.44 | 110.27 | 65.6 | 111.51 | 66.84 | 111.08 | 66.41 | 111.2 | 66.53 | 109.07 | 64.4 | 107.99 | 63.32
| 59191 | 59197 | TS04 | BP21 | BIPM | 59191-59197 | TS04-BP21 | 80.73 | 0.1 | 78.29 | 0.1 | 80.63 | 0.1 | 78.41 |  | 80.58 | 0.1 | 81.89 | 0.1 | 81.58 | 0.1 | 81.62 | 0.1 | 79.57 | 0.1 | 78.31 | 0.1 | TS04-BP21 | 59191-59197 | 27.86 | 43.26 |  | 80.73 | 65.33 | 78.29 | 62.89 | 80.63 | 65.23 | 78.41 | 63.01 | 80.58 | 65.18 | 81.89 | 66.49 | 81.58 | 66.18 | 81.62 | 66.22 | 79.57 | 64.17 | 78.31 | 62.91
Initial run | 59191 | 59197 | TS04 | BP21 | BIPM | 59191-59197 | TS04-BP21 | 80.38 | 0.1 | 77.97 | 0.1 | 80.29 | 0.1 | 78.06 |  | 80.24 | 0.1 | 81.66 | [5 empty] | 79.16 | 0.1 | 77.69 | 0.1
Results for Travel system #2 | [29 empty] | Misclosure |  |  |  | 0.43 |  | 0.53 |  | 0.44 |  | 0.43 |  | 0.42 |  | 0.35 |  | 0.23 |  | 0.31 |  | 0.23 |  | 0.41
|  |  | TS04 | BP21 | BIPM | [22 empty] | TS04-BP21 |  | Mean (chosen value) |  |  |  | 65.545 |  | 63.155 |  | 65.45 |  | 63.225 |  | 65.39 |  | 66.665 |  | 66.295 |  | 66.375 |  | 64.285 |  | 63.115
[30 empty] | Mean (arithmetic) |  |  |  | 65.545 |  | 63.155 |  | 65.45 |  | 63.225 |  | 65.39 |  | 66.665 |  | 66.295 |  | 66.375 |  | 64.285 |  | 63.115
BP1C check
| Start | End | Travel | Local | Labo | Date | Pair | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc |  |  |  |  | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc | Pair | Date | REFDLYT | REFDLYR | Note | P1 (ns) |  | P2 |  | C1
[33 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY | [8 empty] | RAWDIF | DSYSDLY | RAWDIF | DSYSDLY
| 59010 | 59015 | BP1C | BP21 | BIPM | 59010-59015 | BP1C-BP21 | -153.17 | 0.1 | -146.5 | 0.1 | -153.6 | 0.1 | -147.57 | 0.1 | [12 empty] | BP1C-BP21 | 59010-59015 | 234.99 | 72.76 | * | -153.17 | 9.06 | -146.5 | 15.73 | -153.6 | 8.63 | -147.57 | 14.66
| 59191 | 59197 | BP1C | BP21 | BIPM | 59191-59197 | BP1C-BP21 | -182.78 | 0.1 | -176.16 | 0.1 | -183.15 | 0.1 | -177.15 | [13 empty] | BP1C-BP21 | 59191-59197 | 235.71 | 43.26 |  | -182.78 | 9.67 | -176.16 | 16.29 | -183.15 | 9.3 | -177.15 | 15.3
Results for BP1C | [29 empty] | Misclosure |  |  |  | 0.61 |  | 0.56 |  | 0.67
|  |  | BP1C | BP1J | BIPM | [22 empty] | BP1C-BP1J |  | Mean (chosen value) |  |  |  | 9.365 |  | 16.01 |  | 8.965
[30 empty] | Mean (arithmetic) |  |  |  | 9.365 |  | 16.01 |  | 8.965
BP1K check
[8 empty] | RAWDIF(P1) | Unc | RAWDIF(P2) | Unc | RAWDIF(C1) | Unc | RAWDIF(C2) | Unc | RAWDIF(E1) | Unc | RAWDIF(E5) | Unc |  |  |  |  | RAWDIF(B1) | Unc | RAWDIF(B2) | Unc
| 59010 | 59015 | BP1K | BP21 | BIPM | 59010-59015 | BP1K-BP21 | 38.2 | 0.1 | 39.36 | 0.1 | 37.59 | 0.1 | 38.74 | 0.1 | 37.83 | 0.1 | 35.26 | 0.1 | [7 empty] | 0.1 | BP1K-BP21 | 59010-59015 | -32.9 | 72.76 | * | 38.2 | -67.46 | 39.36 | -66.3 | 37.59 | -68.07 | 38.74 | -66.92 | [8 empty] | 0 | -105.66 | 0 | -105.66
| 59191 | 59197 | BP1K | BP21 | BIPM | 59191-59197 | BP1K-BP21 | 8.97 | 0.1 | 10.2 | 0.1 | 8.37 | 0.1 | 9.74 |  | 8.64 | 0.1 | 6.11 | 0.1 | [7 empty] | 0.1 | BP1K-BP21 | 59191-59197 | -32.15 | 43.26 |  | 8.97 | -66.44 | 10.2 | -65.21 | 8.37 | -67.04 | 9.74 | -65.67 | [8 empty] | 0 | -75.41 | 0 | -75.41
Results for BP1K | [29 empty] | Misclosure |  |  |  | 1.02 |  | 1.09 |  | 1.03
|  |  | BP1K | BP21 | BIPM | [22 empty] | BP1K-BP21 |  | Mean (chosen value) |  |  |  | -66.95 |  | -65.755 |  | -67.555
[30 empty] | Mean (arithmetic) |  |  |  | -66.95 |  | -65.755 |  | -67.555