TM155: ftp://tai.bipm.org/TimeLink/LkC/Doc/TM155_BEV_GPS-TW_Calib.doc	               Created 12/08/2008 Last Mod. 5/9/2008

Improving the Calibration of the BEV GPS Receiver Calibrated by Using TWSTFT

Z. Jiang and F. Arias

BIPM

Contact: zjiang@bipm.org

Notations:

TWSTFT	Two Way Satellite Time and Frequency Transfer (TW for short)

AV		GPS All in View time transfer

JRG		Joanneum Research GmbH in Graz

BEV		Bundesamt für Eich-und Vermessungswesen in Vienna

PTB		Physikalisch-Technische Bundesanstalt in Braunschweig

TUG		University of Technology Graz in Graz

Resume: Until recently, the only GPS time receiver calibration technique for the UTC time transfer was the traditional GPS-GPS one [8,9], which has been organised by BIPM for most of the GPS equipment in contributing laboratories. The BEV GPS equipment calibration is the first experiment using the TW-GPS technique and the result has been used for BIPM UTC time transfer. This report studies, in the view of the standard UTC computation, the data processing so as to improve the calibration quality and uncertainty estimation.

1. The JRG calibration result

In cooperation with TUG and with the support of PTB and BEV, using a portable TW ground station, the JRG carried out a TW time transfer between PTB and BEV from 9 to 31 October 2007. The main purpose was to calibrate the BEV GPS time transfer receiver (Rx) through the side by side TW and GPS simultaneous time measurement comparisons [1,10] (cf. Figure 1.1). The reference [10] depicts the equipment setting up for the calibration.

The calibration computation and result were reported in the sections 4.5 and 6 of [1] and [2]. The uncertainty estimation was reported to BIPM by BEV in [3]. The calibration result for the official primary GPS Rx for the UTC time transfer at BEV is:

The above result has been used for the Circular T computation at BIPM.

The calibration value was calculated based on the difference of the TW link of BEV-PTB and the related GPS AV link derived using the Circular T with the following steps:

- Measuring the TW time transfer link between MJD 54383 15h 34m and 54384 13h 55m, totally 121 measured points (Table 24 [1]). Figure 1.2 (left) displays the time deviation of the TW measurements;

- Making a linear regression over the above data set (Figure 1.1 or Diagram 5 in Sect. 4.5 [1]);

- Interpolating the TW time transfer value on MJD 54384 0h using the linear regression obtained. The result is -50.586 ns;

- Deriving, from the Circular T 238 (UTC0710), the GPS AV time transfer value between BEV and PTB, i.e.,  [UTC-UTC(PTB)]-[UTC-UTC(BEV)]= -38.3 ns;

- Comparing TW-GPS link by differing the steps 3. and 4.: -50.586- (-38.3) = -12.3 ns;

- Calculating the internal delay: INT DLY = - 27.3 ns (computation detail was not given in [1,2,3]);

- Finally estimating the calibration uncertainty [3] uB = √[uA(BEV)²+ uA(PTB)²+ uB(PTB)²+ uTW(BEV-PTB)²] = √ [1.5²+0.2²+0.9²+0.782²]=1.9 ns

Figure 1.1 TW time transfer between BEV and PTB (from Tab. 24 of [1]) and the linear regression. See Figure 1.2 left for its TDev. Note: the sign of the BIPM UTC time link is inversed w.r.t. the JRG report.

Figure 1.2 Comparison of the Time Deviations of the links of BEV-PTB: Left TW, Right GPS. The short term time stability of the TW is much better than that of GPS. Here h is hour, d is day and unit is in s-10.

2. Improving the JRG calibration and uncertainty result

We first analyse the JRG result and discuss where and how we can improve it. Then we do the calibration computation and uncertainty estimation. In order to be traceable, we list all the details of the data and the computation procedure. Because the calibration was completely based on the TW and GPS AV time link comparison of the side by side simultaneous measurements, the key is to make an accurate comparison.

2.1 Analysing the JRG calibration

On one side, the linear regression in the step 2 may be affected by the non linear variations and noises in the master clocks’ differences (cf. Figure 1.1) and on the other side, in the step 5, only one point of the Circular T 238 was used. As displayed in the Figure 1.2 (right) and Figure 2.1, the GPS time transfer used for the Circular T 238 was disturbed round about on the MJD on question: 54384. In addition to the bump before 54384, there is unfortunately a gap of about 10 hours in the PTB GPS data file between MJD 54383 18h 36m 30s to MJD 54384 05h 28m 30s, noting that the normal interval is 780 seconds. The value on MJD 54384 of Circular T 238 used for the JRG calibration was without any doubt affected by the disturbance and the gap in the GPS data.

Figure 2.1 GPS AV time transfer link between BEV and PTB used for the Circular T 238. From MJD 54381.1 to 54383.8, the GPS link was disturbed, and just before and after 54384.0, there is a gap of 10 hours. The Circular T 238 was affected by the disturbance. The blue curve is the Vondrak smoothing and the sigma of the residuals of the smoothing is 1.386 ns. The short red line is the 1-day duration of the TW measurements (see Figure 1.1), of which the first half is corresponding to the gap in GPS link data.

2.2 Improving the JRG calibration

A procedure was proposed in [4] to transfer the TW calibration to GPS PPP Rxs. The same procedure can be used to improve the calibration of the BEV GPS MC Rx. As illustrated in Figure 2.1, the second half (MJD 54384.23 - 54384.56) of the 1-day TW measurements covers a period of the non-disturbed GPS measurement, totally 26 GPS points (Tab. 2.2.1) and 41 TW points (Tab. 2.2.2).

Figure 2.2 Comparison of the BEV-PTB time links corresponding to Table 2.2.3. The black crosses are the GPS and the blue squares are the TW. The GPS is much noisy than the TW. The number of the compared points is 26. The mean value between the GPS and TW is -13.983±0.189 ns which is removed in the plot (cf. Tab. 2.2.3).

The principal idea to improve the calibration is to use this common part where both TW and GPS were available and not disturbed to compute the difference between TW and GPS. Table 2.2.3 lists the 26 differences of the two links with the TW measurements interpolated onto the GPS epochs. The GPS data interval is 16 minutes and the TW interval is about 10.5 minutes which is small enough for an accurate interpolation. It is important to note here that, by making the differences between the two links, the clocks are cancelled. The differences are nothing but the total calibration constant plus the measurement errors. The later, as shown in Figures 1.2 and 2.2, mainly come from the GPS measurement instability.  The mean value of the difference is -13.982 ns. The standard deviation of the 26 differences is ±0.964 ns. So the standard deviation of the mean value is ±0.964/√26=±0.198 ns (see table 2.2.3).

2.3 Uncertainty and expanded uncertainty of the improved calibration

The aim of the calibration is to calculate the UTC(BEV)-UTC(PTB). Because the TW link is calibrated, we have UTC(BEV)-UTC(PTB)=TW(BEV-PTB)= GPS(BEV-PTB)-Total Delay. That is, the total delay of the GPS Rx is simply the difference of the TW and GPS links:

Total Delay = GPS(BEV-PTB) - TW(BEV-PTB) = -13.983 ±0.198 ns

Here the ±0.198 ns is the standard deviation of the mean value. The total uncertainty of the TW calibration is 0.9 ns (Table 25 of Section 4.5.3 [1,3,7]). Therefore the total calibration uncertainty of the BEV TTS-2 No. 24 GPS Rx is uB1=√(0.9²+0.198²)=0.92 ns.

The GPS AV data are corrected by using the IGS precise GPS satellite orbits and ionosphere delay information [5]. Because the distance between BEV and PTB is only several hundred kilometres, the biases due to the residual influences can be well cancelled. Even so, because the simultaneous measurements were maintained only 8 hours, the comparison calibration may be subject to unexpected error sources which are not averaged out within 8 hours, such as the multi-paths or diurnal temperature variations or the residual influence of the atmosphere delays etc. On the other side, the long-term instability of the GPS receiver and the related laboratory equipment should be also taken into account in the total uncertainty budget. We hence apply a coverage factor to ensure the uncertainty estimation. If we use the coverage factors 1, 2 and 3, we obtain the uncertainty and the expanded uncertainty estimations:

uB1= 0.92 ns ≈ 1.0 ns (coverage factor = 1)

uB2= 1.84 ns ≈ 2.0 ns (coverage factor =2)

uB3= 2.76 ns ≈ 3.0 ns (coverage factor = 3)

The overage factors 2 and 3 here are to cover the unexpected impacts and the long-term variations in the GPS receiver system. For the factor 3, We have the uB3 =3 ns.

Another estimation can be given based on the hypothesis that the unexpected impacts and the long-term variations are of an order of 2 ns at each side of PTB and BEV. So we have:

uB = √(0.92²+2²+2²) ns = 2.97 ns ≈ 3.0 ns

Both estimations have the same result.

3. Summary

The quality of the calibration is based on the uncertainties of TW link and the comparison of the TW and GPS links. The first is 0.9 ns and the second depends on the methods used. We used the direct comparison of the 26 measurement points and considerably reduced the comparison uncertainty. Doing so, the quality of the calibration is improved.

The calibrated total delay equals the difference of the time links of GPS-TW is -14.3. The internal delay can be derived: INT DLY= - 29.0 ns (instead of -27.3 reported by JRG-BEV [2]). The difference of the calibrations is 1.7 ns. To cover the unexpected impacts and the long-term variations in the GPS receiver system, it is suggested to use the expanded calibration uncertainty uB3 = ±3.0 ns.

Similar calibration can be made for the GPS TTS-2 No 55 if the GPS data are available.

Further improvement of the calibration is possible by increasing the duration of the TW-GPS comparison up to 1-3 days to average out or at least reduce the GPS bias.

Reference

[1] Becker J., Blanzano B., Koudelka O., Merdonig A., Piester D., Ressler H.: Two-Way-Satellite-Time-Transfer between BEV (Vienna) – PTB (Braunscheig), Joanneum Research Report, BEV report, 2007

[2] Niessner A.: BEV Report to BIPM on the result of the calibration of the GPS receivers, 14 Jan. 2008

[3] Niessner A.: email communication to Z. Jiang of BIPM on 28 July and 12 Aug. 2008, see ANNEX

[4] Jiang Z.: Transfer the UTC Calibration from TWSTFT to GPS PPP, BIPM TM 152

[5] Jiang Z. Petit G.: Time transfer with GPS All in View, IN Proc. ATF 2004 pp 236-243

[7] Piester D., Hlavać R., Achkar J., de Jong G., Blanzano B., Ressler H., Becker J., Merck P., Koudelka O.: Calibration of four European TWSTFT earth stations with a portable station through Intelesat 903, IN Proc. EFTF 2005 pp 360-365

[8] De Latour A., Cibiel G., Dantepal J., Dutrey J.-F., Brunet M., Ries L., Issler J.-L.: Dual frequency absolute calibration of GPS receiver for time transfer, IN Proc. EFTF 2005 pp 366-371

[9] Petit G., Jiang Z., Ulrich P., and Taris: Differential calibration of Ashtech Z-12T receiver for accurate time comparisons, IN Proc. EFTF 2000, pp. 40-44, 2000

[10] Piester D, Bauch A, Breakiron L, Matsakis D, Blanzano B, and Koudelka O: Time transfer with nanosecond accuracy for the realization of International Atomic Time, Metrologia, vol. 45, 185 – 198, 2008

ANNEX Emails of the Reference [3]:

(1).

De: Niessner Anton [mailto:Anton.Niessner@bev.gv.at] 
Envoyé: lundi 28 juillet 2008 11:19
À: Zhiheng Jiang
Cc: Elisa Felicitas Arias 
Objet: AW: GPS Calibration Report

Dear Dr. Jiang,

I can inform you that we estimated the uncertainty uB for our reference receiver by rather simple calculation:

uB(BEV) = ((uA(BEV))^2 + (uA(PTB))^2 + (uB(PTB))^2 + (uTWSTFT)^2 )^0.5 =

= (1.5^2 + 0.2^2 + 0.9^2 + 0.782^2 )^0.5 = 1.92653 ns

The specific values are from CircularT and the TWSTFT Calibration Report.

Best regards

Anton Niessner

(2).

De : Niessner Anton [mailto:Anton.Niessner@bev.gv.at] 
Envoyé : mardi 12 août 2008 13:16
À : Zhiheng Jiang
Cc : Elisa Felicitas Arias ; Hawai Konate; Mache Werner
Objet : AW: GPS Calibration Report

Dear Dr. Jiang,

> A question, if I understood well your report, the uncertainty of the TW

> BEV-PTB is 0.9ns ? (Summary in Page 69)

yes, this is the result of the measurement campaign.

> There is a TW working group meeting at SP on 2-3 Oct. Are you going to ?

no, I am not a member of this group.

Regards

Anton Niessner

(3).

De : Zhiheng Jiang [mailto:zjiang@bipm.org]
Envoyé : mercredi 3 septembre 2008 22:23
À : farias@bipm.org
Objet : TM 155

Bonsoir Feli,
> at both ends of the link. E.g. if one considers this long term
instability > to be 2 ns for each end, one obtains sqrt(1+4+4) = 3 ns
If we take the coverage factor 3, we have 0.92*3=2.72 or round off to
uB=3.0 ns as Gérard proposed ?
Jiang

(4)

De : Gérard Petit [mailto:gpetit@bipm.org] Envoyé : mercredi 3 septembre
2008 19:53 À : 'Zhiheng Jiang'; 'Arias'; 'Gianna Panfilo'; 'Hawai Konate';
'Tisserand'; 'Wlodeck'
Objet : RE: TM 155

Bonjour,

>TM 155: File:\\G:\Partage\TM\TM155_BEV_GPS-TW_Caib.doc
<file://SI/COMMUN.Tai/Partage/TM/TM155_BEV_GPS-TW_Caib.doc>
>Comments welcome

The point is, the uB uncertainty is expected to represent the accuracy of
the link at any time during the next year(s). Because there are a number
of factors that can vary in the equipment and environment at both ends of
a link over year(s) (for a GPS link:
multipath, error in troposphere model, error in ionosphere model, effect
of temperature, etc.. even not counting what happens when a connector is
not firmly in place) we have so far used uB = 5 ns as a kind of 'safe value' even though the side by side measurements that constitute the differential calibration itself have an uncertainty well below 1 ns.

So if you have a TW calibration uncertainty at 0.9 ns and a standard
uncertainty of the differential calibration at 0.2 ns (over 8 hours
only!), the uB of the GPS link for these 8 hours may be 1 ns but to
estimate the uB valid for the next year(s) one
has to add an estimation of the long-term instability in the GPS equipment
at both ends of the link. E.g. if one considers this long term instability
to be 2 ns for each end, one obtains sqrt(1+4+4) = 3 ns

Taking uB = 1.8 ns (on the basis of a coverage factor) is equivalent to
estimate the long term instability at each end to be about 0.7 ns which
seems to me a bit optimistic.

In any case, if we take this direction, we should apply it to all
calibrations.

Gérard

# Text boxes

50

45

40

35

30

25

20

54387

54385

54383

54381

54379

54377

38.3 ns

Period of the TW measurements

Tab. 2.2.1 BEV-PTB GPS link /ns

MJD      GPS/ns

54384.22813  37.270

54384.23924  37.300

54384.25035  37.670

54384.27257  38.490

54384.28368  37.210

54384.29479  37.930

54384.30590  38.500

54384.31701  36.310

54384.32813  37.390

54384.33924  38.590

54384.35035  38.500

54384.36146  39.330

54384.37257  35.670

54384.38368  36.230

54384.39479  39.010

54384.40590  38.340

54384.43924  38.320

54384.45035  38.800

54384.46146  38.780

54384.47257  37.670

54384.49479  38.950

54384.50590  37.420

54384.51701  38.180

54384.53924  39.250

54384.55035  39.590

54384.56146  37.830

BEVPTB.MMMA_.dat

Tab. 2.2.3 Differences TW-GPS /ns

Mjd       GPS      TW   GPS-TW

54384.2281  37.270  51.747 -14.477

54384.2392  37.300  51.770 -14.470

54384.2504  37.670  51.758 -14.088

54384.2726  38.490  52.112 -13.622

54384.2837  37.210  51.817 -14.607

54384.2948  37.930  51.629 -13.699

54384.3059  38.500  51.440 -12.940

54384.3170  36.310  51.252 -14.942

54384.3281  37.390  51.028 -13.638

54384.3392  38.590  51.083 -12.493

54384.3504  38.500  51.297 -12.797

54384.3615  39.330  51.623 -12.293

54384.3726  35.670  51.864 -16.194

54384.3837  36.230  52.060 -15.830

54384.3948  39.010  51.866 -12.856

54384.4059  38.340  51.972 -13.632

54384.4392  38.320  52.312 -13.992

54384.4503  38.800  52.709 -13.909

54384.4615  38.780  52.959 -14.179

54384.4726  37.670  53.196 -15.526

54384.4948  38.950  52.192 -13.242

54384.5059  37.420  51.745 -14.325

54384.5170  38.180  51.803 -13.623

54384.5392  39.250  52.916 -13.666

54384.5503  39.590  53.031 -13.441

54384.5615  37.830  52.894 -15.064

N=26,Mean=-13.983,SD=0.964,SD/√N=0.198ns

BEVPTB.GTAT4

Tab. 2.2.2 BEV-PTB TW link /ns

MJD        TW/ns

54384.23194  51.747

54384.24028  51.773

54384.24236  51.912

54384.24444  51.615

54384.24653  51.697

54384.27361  52.129

54384.27986  51.882

54384.31528  51.281

54384.32361  51.139

54384.32569  51.229

54384.32778  51.053

54384.32986  50.903

54384.35694  51.424

54384.36319  51.699

54384.38819  52.140

54384.39028  51.868

54384.39236  51.983

54384.39444  51.844

54384.39861  52.111

54384.40694  51.952

54384.40903  52.363

54384.41111  52.351

54384.41319  52.447

54384.44028  52.307

54384.44653  52.623

54384.47153  53.185

54384.47361  53.207

54384.47569  53.345

54384.47778  52.600

54384.48194  50.768

54384.49028  51.859

54384.49236  51.621

54384.49444  52.291

54384.49653  51.697

54384.52361  51.837

54384.52986  52.819

54384.55486  53.078

54384.55694  53.277

54384.55903  52.811

54384.56111  52.894

BEVPTB.TTTT_.dat [1]

53.5

53.0

52.5

52.0

51.5

51.0

50.5

50.0

54384.55

54384.45

54384.35

54384.25

GPS-TW Link difference = -14.0 ns ± uB

INT DLY= - 29.0 ns ± uB

Expanded uncertainty: uB= ±3.0 ns

_ _ _ _ _ _ _ _ _

(vs. -12.3 and -27.3 ±1.9 ns in the JRG-BEV report [2])

GPS receiver TTS-2 No. 24: INT DLY = - 27.3±1.9 ns

(based on the TW-GPS time link difference: -12.3±1.9ns)

54383.6

54383.8

54384.0

54384.2

54384.4

54384.6

50.586 ns

48

49

50

51

52

53

54384

day

d/2

d/4

d/8

h

h/2

4.3

2.6

3.0

2.5

2.2

day

d/2

d/4

d/8

h

h/2

*

6.2

5.6

7.1

9.7

17.6

TDev / s-10

TW: BEV-PTB

TDev / s-10

GPS: BEV-PTB