TM151_TwCalib.doc,  ftp://tai.bipm.org/TimeLink/LkC/TM151_TwCalib.doc  first 25/5/2008 last 11/7/2008

TWSTFT Calibration Transfer

from UTC Links to Non UTC Links using

Triangle Geometric Condition

Z. Jiang and W. Lewandowski

BIPM

Introduction	2

1. Principle of the TW Calibration	3

2. The Data Set	4

3. The Results of the Triangle Closure Analysis	6

4. The Results of the TW Calibration: CARL and ESDVAR	7

Table 4a Calibrated CALR Values for non UTC links (ESDVAR=NA)	7

Table 4a Calibrated CALR Values for non UTC links (ESDVAR=NA) (continue)	8

Table 4b The Calibrated CALR Values for non UTC links (ESDVAR≠NA)	9

Table 4c CI and uncertainty of the Calibrated Links in Table 4a and 4b	10

5. Estimation of the calibration uncertainty	11

Table 5 Uncertainty estimated (uB/ns)	11

6. Implementation and evaluation of the CALR values	12

6.1 Implementation of the CALR value into the ITU data file	12

6.2 Evaluation of the CALR value for a particular laboratory	12

7. Comparison to the calibrated non UTC links	13

8. Future calibration and re-calibration	14

Acknowledgement	14

References	14

Annexe	15

Summary:

- CALR values are computed for all the European-American TW links;

- The CALR values are listed in Table 4. In Table 4a, the CALR values are computed vs. ESDVAR=0; Table 4b, the CALR values are computed vs. ESDVAR≠0 supplied by the related laboratories;

- The total calibration uncertainties are listed in Table 4c and Table 5;

- Guidelines for implementing and evaluating the CALR values at a laboratory;

- The implementation date is suggested to be MJD 54677 at 0h UTC

Updating history:

- 29/30 June 2008: Table 4a CalR vs. ESDVAR=0; Table 4b CalR vs. ESDVAR≠0 for USNO and NIST

- 3 July 2008: Section 6: Implementation and evaluation of the CALR values; NA ESDVAR=99999.999

- 11 July 2008: Table 4c: CALR CI and uncertainty of the calibrated links

# Introduction

Organising and maintaining the metrological calibration of the time transfer facilities in laboratories contributing to UTC is among the responsibilities of the BIPM.

TWSTFT (Two Way Satellite Time and Frequency Transfer, TW here after) is, together with GPS, the primary time transfer technique used for UTC generation. All the TW links used in the calculation of UTC have been calibrated. The calibration value is listed in the column CALRof the ITU data file. Sometimes a modification of the calibration value is made and the corresponding correction is listed in the column ESDVAR. During the BIPM link computation, both values of CALR and ESDVAR are to be used. However, there are a number of TW observations that are not used in the generation of UTC, some of them providing redundant links, which cannot be used for accurate time transfer since they are not already calibrated. A reason for aiming at calibrating these non UTC un-calibrated TW links is that laboratories could make use of them for time transfer.

On the other hand, as a part of the pilot study of the GPS PPP links [4], the BIPM computes and since April 2008 publishes on its ftp site the link comparison results between GPS PPP and TW. These results are important for the uncertainty studies of both the uA (measurement uncertainty) and the uB (calibration uncertainty) of the two techniques, especially for their long-term variations. Because the TW and GPS time transfer on the same baseline provide redundant links that are independently measured, the comparison between TW and GPS PPP is helpful to study their short and long terms stabilities. Obviously, only calibrated links are useful for monitoring the calibration of the two techniques, and this is a motivation to carry out the calibration of non UTC TW links.

The calibration values listed in Table 4 have been applied to TW links for the calculation of 0804 TW links and GPS PPP experimental links; the link comparison results are published on the BIPM ftp site: ftp://tai.bipm.org/TimeLink/LkC/0804/. Example of the excellent agreement of the traditional calibration for the non UTC TW and GPS PPP links is SP-NIST, for which the mean value of the differences between the two techniques is 0.013 ns with a standard deviation 0.127 ns. Results are available at ftp://tai.bipm.org/TimeLink/LkC/0804/SPNIST/Dlk/SPNIST.TGT35.Gif

The uncertainty of the traditional TW link calibration is uB=1 ns. TW calibration has been made until now in a process that is man-labour, time and money costing and limited by the availability of staff and equipment for calibration. Hence, a more flexible method could allow to easily calibrating all the non UTC links. We propose a method that uses the available redundant TW links. The calibration uncertainty obtained with this method is better than 2 ns (see section 5).

This report presents the method and the results obtained; CALR values are proposed for all the non-UTC TW links. The CALR-ESDVAR values are listed in Table 4a and 4b with all the ESDVAR values set to zero. We propose the laboratories to use these values to update their ITU data files. In the case of non-UTC links with hardware calibration, the results given here can be used as a reference. In fact, the discrepancies between the three calibrations and the triangle closure calibrations are 1.1 ns on average in the given examples (cf. Table 7).

For the newly calibrated TW links, values 99999.999 or 9999999999 that indicate non-availability (NA) in the columns CALR and ESDVAR in the ITU data files should be replaced by the values of CALR and ESDVAR proposed in Table 4. These NA umbers bring problems in the link computation since they are not both compliant with the ITU format. A consequence is the huge CALR or ESDVAR value in the ITU data file. Only the European and American TW laboratories are included in this report.

# 1. Principle of the TW Calibration

We introduce briefly the TW calibration by the triangle closure method.

The principle is very simple: it is possible two calibrate the TW link between two laboratories (LABi and LABj in Fig. 1) if their respective links with the pivot laboratory (PTB in the present context) are calibrated, since they are both UTC links. The calibration is made by making use of the supposing that all TW links with the PTB (that is, UTC links) are calibrated, any non UTC link between two non-pivot laboratories can be calibrated with the so called triangle closure condition: any triangle closure should be equal to zero.

Taken the example of Fig. 1, we intend to calibrate the non UTC link between LABi and LABj. After the calibration, the triangle closure is expected to be zero:

[UTC(LABi)-UTC(PTB)] - [UTC(LABj)-UTC(PTB)]

+ [UTC(LABj)-UTC(LABi)] = Closure → 0				(1)

The non-zero closure should be nothing but the measurement uncertainty. Because the UTC links are calibrated, we have the non-zero-closure = CALR+/-Std, if we set ESDVAR==0. If the ESDVAR values should be kept, for a link with the two end laboratories Lab1 and Lab2, we have, CALR1= - CALR2 = Closure + (ESDVAR2 - ESDVAR1)/2. The final calibration value is the average for a period: typically a UTC month (12 measurements per day over 35 days or 420 measurements for a complete UTC month). To have observations at common dates, values are interpolated at the nearest common TW scheduled epochs and then the closures are calculated. The uncertainty of the calibration can be estimated by the following equation:

u(CloseCalib)=√{ui² + uj²+S²}  						(2)

Here:

ui=u[UTC(LABi)-UTC(PTB)]

uj=u[UTC(LABj)-UTC(PTB)]

S=Std√N where Std is Standard deviation of the triangle closures and N= number of triangles

Values of Std are listed in Table 4. They are used to estimate the calibration uncertainty by using the equation (2). The value S is very small with an order of uA/√N. uA is about 0.5 ns and N is 420 for a UTC monthly data set. The S is then about 0.02 ns. Introducing these values and taking into account of the standard value of uB=1 ns [5,9] and by using equation (2), the uncertainty of this triangle calibration is about 1.4 ns in total, not too much bigger than that of the traditional UTC calibrated links.

# 2. The Data Set

The TAI 0804 (Apr. 2008, Mjd 54554-54584) TW data are used. There are 10 TW laboratories involved. The following figures and tables display the data distribution and availability.

Table 2.1 Laboratories in the UTC time link network of TAI 0804

Fig. 2.1 The UTC time links (blue lines) and the redundant links (black lines)

AOS is available since 0804 and not shown in this figure

Table 2.2 Total independent baselines (Lab2-Lab1)

- Measured = 44, Theory{N(N-1)/2}=45

Table 2.3 Total independent Triangles (Lab1-Lab2-Lab3)

Measured=35; Theory {(N2-3N+2)/2}=36

* Measurements were not available in this calculation

# 3. The Results of the Triangle Closure Analysis

Total and detailed analysis results can be found on the BIPM ftp site:

ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804

Table 3 is an example of the first 20 triangle closures in the triangle PTB-NIST-OP in the file PTB-NIST-OP.Tf2 ( ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804/PTB-NIST-OP.Tf2 ). In the header of the file, we find the details of the statistic results: the number of total measured closures =367, over them all the 3 TW link measurements are available. The standard deviation of the closures is 0.3 ns with the mean value of 7399.7ns. The individual sigma of the 3 links are 0.24ns for NIST-PTB; 0.42ns for OP-PTB and 0.39ns for OP-NIST. Here sigma stands for the estimated uncertainty of each point given by the statistics during the link computation.

Table 3. Example of the triangle closure PTB-NIST-OP0804 (first 20 closures)

Other analysis results can be found at:

The statistics of the triangle closures: ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804/Clos0804.Tf2

Histograms and statistics of the triangle closures:

ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804/Hist0804.Tf2

# 4. The Results of the TW Calibration: CARL and ESDVAR

We propose that starting from TAI0808 (Aug. 2008) TW links which are not used for the computation of UTC provide ITU data files (TwLaboMJ.DDD) where we set initially all the ESDVAR = 0 and the CALR values are those listed in table 4a. In Table 4a and Table 4b, the column No. shows the number of the closure triangle used for the respective CALR computation; column StDev gives the respective standard deviation. Std_N=StDev/√N. N is the triangle number used. The calibration uncertainty is given in next section in Table 5. In the future, only the measured CALR value can replace the value proposed in this table. However, any laboratory can modify the ESDVAR value whenever necessary. For those laboratories, who want to keep their ESDVAR value in this calibration, the related CALR values are listed in Table 4b. The BIPM calibrtation identify code CI and the related uncertainty are listed in Table 4c.

### Table 4a Calibrated CALR Values for non UTC links (ESDVAR=NA)

(unit in ns, Version 29 June 2008)

### Table 4a Calibrated CALR Values for non UTC links (ESDVAR=NA) (continue)

(unit in ns, Version 25 may 2008)

* Remark: according to the ITU format convention, the number 9…9 represents a non-measured value. We set ESDVAR=99999.999 and the related uncertainty ESIG=9.999.

### Table 4b The Calibrated CALR Values for non UTC links (ESDVAR≠NA)

(unit in ns, Version 30 June 2008)

* Remark: according to the ITU format convention, the number 9…9 represents a non-measured value. We set ESDVAR=99999.999 and the related uncertainty ESIG=9.999. For the measured ESDVAR, fill in the measured value and the estimated ESIG.

### Table 4c CALR CI and uncertainty of the Calibrated Links in Table 4a and 4b

| CALR CI | LOC-REM | S | CALR local /ns | CALR remote /ns | TYPE | Uncert. /ns | Used since | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 301 | AOS/CH | 1 | 21,039 | -21,039 | T.C.* | 6.0 | 30 July 2008 | TM151 |
| 302 | AOS/IT | 1 | 132,777 | -132,777 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 303 | AOS/OP | 1 | 7133,267 | -7133,267 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 304 | AOS/ROA | 1 | 105,365 | -105,365 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 305 | AOS/SP | 1 | 4,478 | -4,478 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 306 | AOS/VSL | 1 | 117,862 | -117,862 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 307 | CH/IT | 1 | 110,859 | -110,859 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 308 | CH/OP | 1 | 7112,132 | -7112,132 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 309 | CH/ROA | 1 | 83,825 | -83,825 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 310 | CH/SP | 1 | -16,414 | 16,414 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 311 | CH/VSL | 1 | 96,498 | -96,498 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 312 | IT/OP | 1 | 7000,408 | -7000,408 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 313 | IT/ROA | 1 | -27,984 | 27,984 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 314 | IT/SP | 1 | -128,372 | 128,372 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 315 | IT/VSL | 1 | -14,696 | 14,696 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 316 | OP/ROA | 1 | -7028,065 | 7028,065 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 317 | OP/SP | 1 | -7128,663 | 7128,663 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 318 | OP/VSL | 1 | -7015,433 | 7015,433 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 319 | ROA/SP | 1 | -100,440 | 100,440 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 320 | ROA/VSL | 1 | 12,746 | -12,746 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 321 | SP/VSL | 1 | 113,117 | -113,117 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 322 | NIST/AOS | 1 | 154,480 | -154,480 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 323 | USNO/AOS | 1 | 403,432 | -403,432 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 324 | NIST/CH | 1 | 176,060 | -176,060 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 325 | USNO/CH | 1 | 425,057 | -425,057 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 326 | NIST/IT | 1 | 285,833 | -285,833 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 327 | USNO/IT | 1 | 534,735 | -534,750 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 328 | NIST/OP | 1 | 7287,687 | -7287,687 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 329 | NIST/ROA | 1 | 258,436 | -258,436 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 330 | NIST/SP | 1 | 159,322 | -159,322 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 331 | NIST/VSL | 1 | 273,323 | -273,323 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 332 | OP/USNO | 1 | 7536,583 | -7536,583 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 333 | USNO/ROA | 1 | 507,564 | -507,564 | T.C. | 6.0 | 30 July 2008 | TM151 |
| 334 | USNO/SP | 1 | 408,247 | -408,247 | T.C. | 2.0 | 30 July 2008 | TM151 |
| 335 | USNO/VSL | 1 | 522,444 | -522,444 | T.C. | 2.0 | 30 July 2008 | TM151 |

* T.C.=Triangle Closure

# 5. Estimation of the calibration uncertainty

We present hereafter only a general description of the uncertainty estimation listed in Table 4c. A more detailed discussion will be given in a separate paper.

From Section 6 of BIPM Circular T 244 [5] and [9], the traditional calibration uncertainty uB is 1ns for TW links and 5ns for GPS links. In table 4, the maximum Std is 1.6 and the minimum N is 169. In the worst case, the S equals 0.12ns. Introducing them to equation (2), we have:

for TW-GPS mixed triangles: uB1 = √{5² + 1²+ 0.12²} = 5.1 ns  < 6ns

for TW triangles: uB2 = √{1² + 1²+ 0.12²} = 1.4 ns < 2ns

We have thus set the calibration uncertainties to:

### Table 5 Uncertainty estimated (uB/ns)

| Labs/Links | uB |
| --- | --- |
| AOS, NIST, ROA concerned links (with PTB, USNO, OP, IT, VSL, CH and SP) | 6 |
| all the other links in Table 4a and 4b | 2 |

# 6. Implementation and evaluation of the CALR values

The implementation date is suggested to be MJD 54677 at 0h UTC.  TW labs are suggested to make available the data as often as possible so that all the labs and BIPM can make parallel evaluations. Data exchanges can be done through the usual BIPM UTC ftp uploading site (see bellow).

## 6.1 Implementation of the CALR value into the ITU data file

Some details when implementing the CalR value in the ITU files:

  - For those use the Table 4a, the ESDVAR should be 99999.999 and related ESIG=9.999;

  - For those use the Table 4b, set the real value for ESDVAR and ESIG;

  - Inserted the following line in the header of the ITU file:

CALR CI TYPE: TRIANGLE CLOSURE MJD: 54554-54584 EST. UNCERT.: x.x ns

Here the CI is the BIPM calibration identifier series number and x.x is the uncertainty of the calibration given in Table 4c (2 or 6 ns).

To easy data exchanges between the labs for the mutual evaluations, it is suggested to regularly upload and download the ITU data files through the BIPM ftp site:

URL address	: tai.bipm.org

IP address	: 62.161.69.131

FTP:

Port		: 21

User		: labotai

Password	: dataTAI

A specific directory for data of each laboratory was created. You can put your data in the remote directory "/data/Lab/links/TWSTFT/" where "Lab" is the BIPM acronym of your laboratory.

## 6.2 Evaluation of the CALR value for a particular laboratory

The aim of the evaluation is to find the unexpected errors in the CALR values in Table 4. A laboratory can use the following steps to evaluate its CALR values for a non-UTC link with the help of a BIPM Circular T:

- Use a recent Circular T and the TW data after MJD 54584 so that the raw data are independent from that used for the calibration in Table 4 of this TM; this will provide a means of  checking Table 4;

- Implement the CALR values in Table 4a or 4b into the ITU TW data file. Carefully respect the ITU format convention. For reference, you find your ESDVAR and CALR values (TAI0804) in ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804/T314.OUT.@;

- Use the Circular T to compute the indirect links. For example for NIST and IT, the indirect link via PTB is NIST-IT=-(IT-NIST)=(NIST-PTB)-(IT-PTB)={(UTC-UTC(IT)-[UTC-UTC(NIST)]}. Here the pivot lab PTB is cancelled;

- Compute the direct link NIST-IT using the ITU data;

- Interpolate the direct link values onto the standard UTC MJDs;

- Compute the difference between the indirect link (3) and the direct link (4,5);

- Re-do 3-5 but in the opposite order, you should have IT-NIST=-(NIST-IT);

- Above is the evaluation over the standard MJDs with the help of Circular T. One can also compute the closures for the triangles on all the measurement time epochs (12 points per day). It is expected that the closure≈0.  However the above steps 1-7 should be enough for the evaluation. In fact, BIPM has made a global closure analysis for the data set TAI0804: ftp://tai.bipm.org/TFG/JIANG/NetWork/TW_Calib0804/ ;

- Analyse the differences or the closures which should be smaller or equal to the tolerance: √[uA1²+ uA2²+ uA3²+S²] (1-σ).

Here uA1, uA2 and uA3 are the link uncertainties of the three links concerned. Approximately, we can use the uA listed in the Section 6 of the Circular T, uA=0.5ns. S is the precision of CALR, see equation (2), which is very small. Therefore the tolerance of the differences is about 0.9 ns (1-σ).

Therefore, if a difference is bigger than 1 ns, further investigation should be made to know what to identify the causes.  It can be due to:

- Data missing: there are gaps in the raw data;

- Diurnals, especially for the Europe-America links. The amplitude may be up to 1 ns or bigger;

- Outliers;

- Different calibration used for Circular T and the Tab. 4;

- Clock noises and interpolation errors: if differences are within 2 ns but there is not a systematic bias (do not keep the same sign), it is acceptable;

- If a bias bigger than 2 ns is present, then verify all the old calibrations. If no clear reason is found, it is suggested to accept the Table 4 values that agree with the official UTC link calibration;

- the uB proposed for the CALR is 2 ns or 6 ns depending on the original calibration (Tab 4c).

# 7. Comparison to the calibrated non UTC links

As a support to the change of satellite for the TW observations in February 2008, the BIPM computed the GPS carrier phase bridges [6,7,8] for the UTC and non UTC links. Among them there are three OP concerned non-UTC links which had been classically calibrated and the calibration values are not applied in the Lab1(OP) ITU data files. They can be used to check the quality of our triangle closure calibrations in Table 4. As an effect of the bridge calibration, the uncertainty should slightly increase to about 1.5ns. We have then compared the traditional calibration and our triangle closure calibration. The results are in Table 7:

Table 7 Differences between traditional calibration and the triangle closure calibration (ns)

| Link | Bridge | CALR (Tab.4) | Difference | Comments |
| --- | --- | --- | --- | --- |
| SP-OP | 7127.8 | 7128.7 | -1.1 |  |
| VSL-OP | 2014.5 | 2015.4 | -0.9 | CALR-5000 |
| IT-OP | -1.0 | 0.4 | -1.4 | CALR-7000 |

The differences are well below the total uncertainty: √{1.5² + 2²} = 2.5 ns. This agrees well with the proposed CALR values given in Table 4 and the uncertainty estimation of 2ns. However, if we take a closer look at the differences: they are all negative round about the averaged value -1.1 ns. This implies, the discrepancies here come from rather the systematic bias of the traditional calibrations or the bridges than from this triangle closure calibration. The latter is dominated by measurement uncertainty uA which obeys the white noise ruler.

# 8. Future calibration and re-calibration

Four different cases should be considered:

- Change of satellites as happened in February 2008: because all the TW laboratories are backed up by GPS PPP (except for AOS), it is suggested to use the GPS PPP solutions to bridge the UTC links and then use the UTC links to calibrate the non UTC links like what we do in this report;

- In case of having some TW links re-calibrated, the same method should be applied to re-calibrate the non UTC links;

- Asian Laboratories;

- New TW laboratories: TUG, IPQ etc.

Because the real uncertainties of the traditional and the triangle closure calibrations are not very different and both are sufficient for usual clock comparisons, it is suggested to perform the non UTC link calibration only when and where necessary.

BIPM is actively contributing to the organization and maintenance of the metrological calibration of the UTC time transfer facilities and has established a procedure in Tsoft to perform these calibrations (cf. Annexe).

# Acknowledgement

We thank Dr. Arias, Dr. Piester, Dr. Matsakis, Dr. Zhang and Dr. Lorini for their carefully studying this report and helpful comments. We thank all the TW laboratories for their active cooperation.

# References

[1] Jiang Z., D. Matsakis and W. Lewandowski: TW network time transfer and baseline dependent uncertainties analysis, TM 138, 5 May 2005

[2] Jiang Z.: From TW Single Baseline to Network Time Transfer, TM 138bis, Apr. 2008

[3] Bauch A., L. Breakiron, D. Matsakis, D. Piester, Z; Jiang: On the accuracy of Two-Way satellite time and frequency transfer: a study of triplet closures, reported to CPEM 2008

[4] BIPM pilot study in GPS PPP for TAI time transfer

[5] BIPM Circular T 244, Section 6, May 2008

[6] Piester D, D. Matsakis, L. Lorini, V. Zhang.: email communications, Apr.-July 2008

[7] BIPM Circular T 242, Feb. 2008

[8] BIPM Report for Circular T 242, Feb. 2008

[9] Piester D., A. Bauch, L. Breakiron, D. Matsakis, B. Blanzano, O. Koudelka: Time transfer with nanosecond accuracy for the realization of International Atomic Time, Metrologia, vol. 45, 185 – 198, 2008

# Annexe

This section is only for the BIPM time section staff:

We use the Tsoft/F2/F3162 to calculate the calibration following the steps below:

- run1: Y4   - UTC links with PTB with the options ON: CALR  EDSVAR

- run2: Y4   - Non UTC links (No PTB) with the options OFF: -CALR  -EDSVAR

- run3: F2    - CalR by triangle method, better with the option ON:  LKTY4i (Lab1==PTB)

- run4: Y20 - Set CalR calculated/ESDVAR==0 into TwYYMM.lk (using this file)

- run5: Y4   - UTC+Non UTC links with the option ON:  CALR  EDSVAR

- run6: F2    - Verify if all the triangle closures -- >0

The calibration results are in Table 4 TW Calibrations of CALR and ESDVAR. Certainly, if any of the non-UTC links are hardware calibrated, the values here can be used as a check (cf. section 6).

(Job for F2: 0804 LKTY4i -LKGY3i -ALIGN -ALIGN.G2T -ALIGN.T2G -Out888 OutStep=1 HistScale=1000 HistIntv=250 MjdCol01_09 LinkCol18_30)

If a laboratory wants to keep its ESDVAR value, this value should be subtracted from the originally determined CalR values (X:\TaN\0804\Adj\CalR0804.Tf2) which is computed by F2 supposing EsdVar==0. The new CalR value with EsdVar-removed can be computed by:

CalR1(new)= -CalR2(new)={CalR1(old) + ( [EsdVar2-EsdVar1]/2 ) }

Use Y20 option: RmEsdVar ON to make an iteration to remove EsdVar from the original file X:\TaN\0804\Adj\CalR0804.Tf2 and then re-establish the big TW file for TAI computation:

- run1: Type in Link_LAB1_Non-Zero_EsdVar value in X:\TaN\0804\Adj\CalR0804.Tf2 (note here for a link in CalR0804.Tf2 type in the Lab1 value. For example:

Calib. USNO01   CH01: S=1 CALR=   425.057 +/- 0.426, ESDVAR=  -387.250  375 0.022   14

- run2: Y20 with Calib & RmEsdVar ON. This will read the above copied file in which the ESdVar(sb) should be removed from the listed CalR values: Output new files is X:\TaN\0804\Adj\CalR0804.Tf2.Y20

- run3: Copy the NewCalR file: X:\TaN\0804\Adj\CalR0804.Tf2.Y20 (present only the non-zero-EsdVar labs) into the old: X:\TaN\0804\Adj\CalR0804.Tf2

- run4: Re-do Y20 with Calib ON and RmEsdVar OFF

# Text boxes

IT

Labi Labj   S  CARL                StDev  ESDVAR*             N  Std_N

IT   OP: S=1 CALR=  7000.408 +/- 0.185, ESDVAR= 99999.999  343 0.010

OP   IT: S=1 CALR= -7000.408 +/- 0.185, ESDVAR= 99999.999  343 0.010

IT  ROA: S=1 CALR=   -27.984 +/- 0.246, ESDVAR= 99999.999  292 0.014

ROA   IT: S=1 CALR=    27.984 +/- 0.246, ESDVAR= 99999.999  292 0.014

IT   SP: S=1 CALR=  -128.372 +/- 0.232, ESDVAR= 99999.999  359 0.012

SP   IT: S=1 CALR=   128.372 +/- 0.232, ESDVAR= 99999.999  359 0.012

IT  VSL: S=1 CALR=   -14.696 +/- 0.998, ESDVAR= 99999.999  236 0.065

VSL   IT: S=1 CALR=    14.696 +/- 0.998, ESDVAR= 99999.999  236 0.065

OP  ROA: S=1 CALR= -7028.065 +/- 0.271, ESDVAR= 99999.999  297 0.016

ROA   OP: S=1 CALR=  7028.065 +/- 0.271, ESDVAR= 99999.999  297 0.016

OP   SP: S=1 CALR= -7128.663 +/- 0.245, ESDVAR= 99999.999  365 0.013

SP   OP: S=1 CALR=  7128.663 +/- 0.245, ESDVAR= 99999.999  365 0.013

OP  VSL: S=1 CALR= -7015.433 +/- 1.245, ESDVAR= 99999.999  232 0.082

VSL   OP: S=1 CALR=  7015.433 +/- 1.245, ESDVAR= 99999.999  232 0.082

ROA   SP: S=1 CALR=  -100.440 +/- 0.307, ESDVAR= 99999.999  314 0.017

SP  ROA: S=1 CALR=   100.440 +/- 0.307, ESDVAR= 99999.999  314 0.017

ROA  VSL: S=1 CALR=    12.746 +/- 1.305, ESDVAR= 99999.999  217 0.089

VSL  ROA: S=1 CALR=   -12.746 +/- 1.305, ESDVAR= 99999.999  217 0.089

SP  VSL: S=1 CALR=   113.117 +/- 1.361, ESDVAR= 99999.999  245 0.087

VSL   SP: S=1 CALR=  -113.117 +/- 1.361, ESDVAR= 99999.999  245 0.087

END --- F3162: CALR by triangle differential calibration

1.  CH   IT

2.  PTB  CH

3.  IT   PTB

4.  CH   NIST

5.  NIST IT

6.  USNO CH

7.  NIST USNO*

8.  CH   OP

9.  VSL  CH

10.  OP   VSL

11.  SP   CH

12.  PTB  SP

13.  CH   ROA

14.  ROA  OP

15.  SP   USNO

16.  USNO ROA

17.  VSL  NIST

18.  USNO IT

19.  IT   OP

20.  VSL  IT

21.  SP   IT

22.  IT   ROA

23.  PTB  NIST

24.  NIST OP

25.  SP   NIST

26.  NIST ROA

27.  PTB  OP

28.  USNO OP

29.  SP   OP

30.  USNO PTB

31.  VSL  PTB

32.  PTB  ROA

33.  VSL  ROA

34.  SP   ROA

35.  VSL  SP

36.  VSL  USNO

1 NIST 10002

2 PTB  10005

3 USNO 10007

4 OP   10008

5 IT   10011

6 ROA  10014

7 VSL  10024

8 CH   10057

9 SP   10072

10 AOS  10046

Tsoft/F3162/F3120: CALR differential calibrationby the Triangle closure method*/

Remove ESDVARs from CALR: CalR1=-CalR2=CalR1+[EsdVar2-EsdVar1]/2

Labi Labj   S  CARL                StDev  ESDVAR*              N Std_N

NIST  AOS: S=1 CALR=   154.480 +/- 0.335, ESDVAR=   224.040  326 0.019

AOS NIST: S=1 CALR=  -154.480 +/- 0.335, ESDVAR= 99999.999  326 0.019

USNO  AOS: S=1 CALR=   403.432 +/- 0.286, ESDVAR=  -387.250  277 0.017

AOS USNO: S=1 CALR=  -403.432 +/- 0.286, ESDVAR= 99999.999  277 0.017

NIST   CH: S=1 CALR=   176.060 +/- 0.420, ESDVAR=   224.040  384 0.021

CH NIST: S=1 CALR=  -176.060 +/- 0.420, ESDVAR= 99999.999  384 0.021

USNO   CH: S=1 CALR=   425.057 +/- 0.426, ESDVAR=  -387.250  375 0.022

CH USNO: S=1 CALR=  -425.057 +/- 0.426, ESDVAR= 99999.999  375 0.022

NIST   IT: S=1 CALR=   285.833 +/- 0.398, ESDVAR=   224.040  360 0.021

IT NIST: S=1 CALR=  -285.833 +/- 0.398, ESDVAR= 99999.999  360 0.021

USNO   IT: S=1 CALR=   534.735 +/- 0.336, ESDVAR=  -387.250  351 0.018

IT USNO: S=1 CALR=  -534.735 +/- 0.336, ESDVAR= 99999.999  351 0.018

NIST   OP: S=1 CALR=  7287.687 +/- 0.292, ESDVAR=   224.040  367 0.015

OP NIST: S=1 CALR= -7287.687 +/- 0.292, ESDVAR= 99999.999  367 0.015

NIST  ROA: S=1 CALR=   258.436 +/- 0.437, ESDVAR=   224.040  316 0.025

ROA NIST: S=1 CALR=  -258.436 +/- 0.437, ESDVAR= 99999.999  316 0.025

NIST   SP: S=1 CALR=   159.322 +/- 0.495, ESDVAR=   224.040  384 0.025

SP NIST: S=1 CALR=  -159.322 +/- 0.495, ESDVAR= 99999.999  384 0.025

NIST  VSL: S=1 CALR=   273.323 +/- 1.569, ESDVAR=   224.040  246 0.100

VSL NIST: S=1 CALR=  -273.323 +/- 1.569, ESDVAR= 99999.999  246 0.100

USNO   OP: S=1 CALR=  7536.583 +/- 0.311, ESDVAR=  -387.250  358 0.016

OP USNO: S=1 CALR= -7536.583 +/- 0.311, ESDVAR= 99999.999  358 0.016

USNO  ROA: S=1 CALR=   507.564 +/- 0.440, ESDVAR=  -387.250  306 0.025

ROA USNO: S=1 CALR=  -507.564 +/- 0.440, ESDVAR= 99999.999  306 0.025

USNO   SP: S=1 CALR=   408.247 +/- 0.515, ESDVAR=  -387.250  374 0.027

SP USNO: S=1 CALR=  -408.247 +/- 0.515, ESDVAR= 99999.999  374 0.027

USNO  VSL: S=1 CALR=   522.444 +/- 1.547, ESDVAR=  -387.250  235 0.101

VSL USNO: S=1 CALR=  -522.444 +/- 1.547, ESDVAR= 99999.999  235 0.101

END --- F3120: Remove EsdVar from CalR. Re-Run Y20 after coping the above table into X:\TaN\0804\Adj\CalR0804.Tf2

LABj-PTB

Fig. 1 Non UTC TW link calibration by triangle closures

CALR(LABj-LABi)=[UTC(LABi)-UTC(PTB)]-[UTC(LABj)-UTC(PTB)]+[UTC(LABj)-UTC(LABi)]

LABj-PTB

LABi- LABj

PTB

LABj

LABi

PTB-NIST-OP0804.Tf2

Tai0804 Closure: T PTB NIST OP(21+13+32)/02_05 05_08 08_02/

N,Min,Max,Mean,RMS,Std: 367 7398.9 7400.6 7399.7 7399.7 0.3

Sigma of the three concerned links: +- 0.24 0.42 0.39

MJD       Closure    Link1   Link2     Link3

1 54554.033   7399.684 = 34.075 -16.691 +7382.300

2 54554.116   7399.468 = 34.008 -16.836 +7382.296

3 54554.199   7399.372 = 33.894 -17.733 +7383.211

4 54554.283   7399.765 = 33.922 -17.552 +7383.395

5 54554.366   7399.507 = 34.191 -16.781 +7382.097

6 54554.449   7399.623 = 34.101 -16.467 +7381.989

7 54554.533   7399.059 = 33.846 -17.009 +7382.222

8 54554.616   7399.265 = 34.130 -17.325 +7382.460

9 54554.699   7399.781 = 34.411 -17.103 +7382.473

10 54554.783   7400.049 = 34.434 -17.390 +7383.005

11 54554.866   7399.901 = 34.366 -16.819 +7382.354

12 54554.949   7399.642 = 34.364 -16.990 +7382.268

13 54555.033   7399.294 = 34.148 -16.821 +7381.967

14 54555.116   7399.106 = 34.157 -17.496 +7382.445

15 54555.199   7399.362 = 34.232 -18.012 +7383.142

16 54555.283   7399.446 = 34.317 -18.234 +7383.363

17 54555.366   7399.911 = 35.115 -18.659 +7383.455

18 54555.449   7399.675 = 34.956 -18.321 +7383.040

19 54555.533   7399.342 = 34.238 -18.145 +7383.249

20 54555.616   7399.189 = 34.315 -18.988 +7383.862

… …

Tsoft/F2/F3162: CALR differential calibration by the Triangle closure method*

Labi Labj   S  CARL                StDev  ESDVAR*              N Std_N

AOS   CH: S=1 CALR=    21.039 +/- 0.227, ESDVAR= 99999.999  288 0.013

CH  AOS: S=1 CALR=   -21.039 +/- 0.227, ESDVAR= 99999.999  288 0.013

AOS   IT: S=1 CALR=   132.777 +/- 0.313, ESDVAR= 99999.999  344 0.017

IT  AOS: S=1 CALR=  -132.777 +/- 0.313, ESDVAR= 99999.999  344 0.017

AOS   OP: S=1 CALR=  7133.267 +/- 0.320, ESDVAR= 99999.999  329 0.018

OP  AOS: S=1 CALR= -7133.267 +/- 0.320, ESDVAR= 99999.999  329 0.018

AOS  ROA: S=1 CALR=   105.365 +/- 0.400, ESDVAR= 99999.999  241 0.026

ROA  AOS: S=1 CALR=  -105.365 +/- 0.400, ESDVAR= 99999.999  241 0.026

AOS   SP: S=1 CALR=     4.478 +/- 0.533, ESDVAR= 99999.999  348 0.029

SP  AOS: S=1 CALR=    -4.478 +/- 0.533, ESDVAR= 99999.999  348 0.029

AOS  VSL: S=1 CALR=   117.862 +/- 0.478, ESDVAR= 99999.999  231 0.031

VSL  AOS: S=1 CALR=  -117.862 +/- 0.478, ESDVAR= 99999.999  231 0.031

CH   IT: S=1 CALR=   110.859 +/- 0.484, ESDVAR= 99999.999  357 0.026

IT   CH: S=1 CALR=  -110.859 +/- 0.484, ESDVAR= 99999.999  357 0.026

CH   OP: S=1 CALR=  7112.132 +/- 0.513, ESDVAR= 99999.999  367 0.027

OP   CH: S=1 CALR= -7112.132 +/- 0.513, ESDVAR= 99999.999  367 0.027

CH  ROA: S=1 CALR=    83.825 +/- 0.585, ESDVAR= 99999.999  172 0.045

ROA   CH: S=1 CALR=   -83.825 +/- 0.585, ESDVAR= 99999.999  172 0.045

CH   SP: S=1 CALR=   -16.414 +/- 0.625, ESDVAR= 99999.999  377 0.032

SP   CH: S=1 CALR=    16.414 +/- 0.625, ESDVAR= 99999.999  377 0.032

CH  VSL: S=1 CALR=    96.498 +/- 1.130, ESDVAR= 99999.999  245 0.072

VSL   CH: S=1 CALR=   -96.498 +/- 1.130, ESDVAR= 99999.999  245 0.072

T PTB AOS IT

T PTB AOS NIST

T PTB AOS OP

T PTB AOS ROA

T PTB AOS SP

T PTB AOS USNO

T PTB AOS VSL

T PTB CH IT

T PTB CH NIST

T PTB CH OP

T PTB CH ROA

T PTB CH SP

T PTB CH USNO

T PTB CH VSL

T PTB IT NIST

T PTB IT OP

T PTB IT ROA

T PTB IT SP

T PTB IT USNO

T PTB IT VSL

T PTB NIST OP

T PTB NIST ROA

T PTB NIST SP

T PTB NIST USNO*

T PTB NIST VSL

T PTB OP ROA

T PTB OP SP

T PTB OP USNO

T PTB OP VSL

T PTB ROA SP

T PTB ROA USNO

T PTB ROA VSL

T PTB SP USNO

T PTB SP VSL

T PTB USNO VSL

T PTB AOS CH

SP

NIST

USNO

ROA

VSL

OP

PTB

CH

37.  AOS PTB

38.  AOS CH

39.  AOS IT

40.  AOS NIST

41.  AOS OP

42.  AOS ROA

43.  AOS SP

44.  AOS USNO

45.  AOS VSL