BIPM TM176 (ftp://tai.bipm.org/TimeLink/LkC/)								version 23 Oct. 2010

BIPM Calibrations of AOS, IPQ, NPL and USNO TWSTFT Time Links

Z. Jiang, W. Lewandowski, A. Harmegnies and D. Piester

1. Introduction	2

2. Implementation schedule	2

Remark	2

Table 1: Alignment of the TW KU UTC link to TW X or GPS MC	3

Figure 1. Alignment of TW to GPS MC for the baseline IPQ-PTB	3

3. Results	4

Table 2 CALR values of the UTC Links	4

Table 3a CALR/ESDVAR for UTC links	4

Table 3b CALR for Non-UTC links	5

4. Discussions	6

4.1 Statistics of the triangle closures	6

Figure 4.1 the histogram of the triangle closures of PTB, CH and USNO	6

Table 4. Verification of the new CALR values with the Triangle closures data of TW1008 and 1009 MJD 55432-55469	7

4.2 Comparison of the TW and GPS PPP UTC time links	8

Table 5 Differences between TW and GPS over the UTC baselines	8

4.3 Comparison of the TW and GPS PPP Non UTC time links	8

5. References:	8

Annex 1. How to make the Implementation	9

A1.1 ITU data file header entries	9

A1.2 Examples for ITU-format data recordings	10

Annex 2. Link comparisons on the baseline USNO-PTB	11

Annex 3. Link comparisons on the UTC baselines	12

Annex 4. Link comparisons on the calibrated Non UTC baselines	16

Annex 5. Triangle closures of the BIPM 2009 and 2010 calibrated time links	26

Implementation schedule:

- Between 24-30 Oct. 2010: The TW LABs check their own CALR and the related ESDVAR in Table 3a and 3b

- The implementation date is at 0h 55502, 2 Nov. 2010: The TW labs implement the new CALR/ESDVAR and the CI codes in their ITU files (c.f. examples in Annex 1). You can do it during the whole month of Nov. and then upload (re-upload) the corrected ITU files to the BIPM ftp site for the UTC computation. The first MJD of the next Circular T 275 is 55504. The computation of CirT 275 will be performed in the earlier Dec.

## 1. Introduction

The last “BIPM” calibration was the Restoration of the TWSTFT link calibration using GPS bridging after the satellite switch from IS-3R to T-11N in July/August 2009, carried out in Aug. 2009 and implemented in 55102-55108 (28 Sept - 4 Oct), cf. BIPM TM 166 [1]. After several successful “BIPM” calibrations, a standard procedure is established at BIPM so as to supply these kind of service when necessary, cf. [2] for details. This includes mainly the following operations:

- Restore and/or align the TW to a calibrated link, in most case to GPS

- Compute the CALR values and uncertainties for the UTC links (LABj-PTB), taking into account of the related ESDVAR values

- Compute the CALR values and uncertainties for the Non UTC links (LABj-LABi) with the method TCC (triangle closure calibration), taking into account of the related ESDVAR values

- Assigning the CI code to the new calibration

- Preparing the calibration report (TM)

- Verifications of the results

- The TW Labs implement the new CALR/ESDVAR values and the related CI codes in the ITU data files

- The TW Labs upload/re-upload the corrected ITU files to the BIPM ftp site for the next circular T computation

Due to operation/equipment failures, NPL and AOS did not take part in the last “BIPM” calibration in Aug. 2009. Then NPL and AOS came back separately in the end of 2009 and in Oct. 2010. IPQ participated in UTC generation in July 2010. On the other side, studies showed that the USNO-PTB TW KU link suffered a jump about 3 ns with respected to the TW X band and GPS PPP (cf. Annex 2), it was decided to apply a link correction of 3 ns in three steps of 1 ns each. These steps were performed in Jan., Feb. and March 2010. All these modifications require a new BIPM calibration.

This TM176 presents the results of the calibration which concerns only the links with the 4 labs: AOS, NPL, IPQ and USNO. We give here directly results without discussing the methodology. Please cf. [1,2] for the method used in details.

## 2. Implementation schedule

- Between the 24-30 Oct. 2010: The TW LABs check their own CALR and the related ESDVAR in Table 3a and 3b

- The implementation date is dated on the 0h 55502, the 2 Nov. 2010: The TW labs implement the new CALR/ESDVAR and the CI codes in their ITU files (c.f. examples in Annex 1). You can do it during the whole month Nov. and then upload (re-upload) the corrected ITU files to the BIPM ftp site for the UTC computation. The first MJD of the next Circular T 275 is 55504. The computation of CirT 275 will be performed in the earlier Dec.

### Remark

  - We used the GPS MC to make the alignment because the P3 receivers are not calibrated

  - Due to the 1009 data uncompleted, we used one and half month’s data from Mjd 55420-55469

  - Annex 2 to Annex 5 are mainly for the BIPM documentation of this work. People, who are not interesting in details of the computations, can ignore these annexes.

Table 1 presents the results of the alignment computations. We tried to use all the available data. However, the only available data of the baseline IPQ-PTB were from MJD 55376 to 55387 (Figure 1). The GPS MC and P3 before and after this period were extremely disturbed and were not applicable. As illustrated in the Annex 2, the USNO-PTB KU band link was calibrated through the X band in the end of 2008 by USNO (personal communication with D. Matsakis) and then drifted about 2-3 ns during 2009. Earlier studies showed that this drift was stabled with respect to the X band on about Nov. 2009 after, when the difference was -3.5 ns. The recomputed D of NPL-PTB is almost the same as that of 1 year ago.

### Table 1: Alignment of the TW KU UTC link to TW X or GPS MC

| Baseline | Aligned to | Common Points | D±StDev | Remark |
| --- | --- | --- | --- | --- |
| USNO-PTB | Xband | 340 | -3.5±0.9 | 0911 |
| AOS-PTB | MC | 263 | 165.0±1.3 | 1008/1009 |
| IPQ-PTB | MC | 131 | 188.6±2.2 | 1007 |
| NPL-PTB | MC | 410 | -490.3±0.9 | 1008/1009 |

### Figure 1 Alignment of TW to GPS MC for the baseline IPQ-PTB

## 3. Results

### Table 2 CALR values of the UTC Links

| LABi–LABj | CALR (ns) | CI | S | Alignment Uncertainty (ns) | Uncertainty to be reported in ITU-file header (ns) |
| --- | --- | --- | --- | --- | --- |
| USNO-PTB | 214.5 | 174 | 1 | 0.2 | 3.0 |
| AOS-PTB | -165.0 | 171 | 1 | 0.2 | 7.0 |
| IPQ-PTB | -188.6 | 172 | 1 | 0.4 | 7.0 |
| NPL-PTB | 490.3 | 173 | 1 | 0.2 | 7.0 |

### Table 3a CALR/ESDVAR for UTC links

### Table 3b CALR for Non-UTC links

Remarks: there were almost not data for the NPL concerned links with AOS, CH, NIST, USNO and VSL. These links are not calibrated in this computation.

## 4. Discussions

In this section, we will make some detailed discussions which are rather for the BIPM technical back up and documentation of this work. This may be helpful for those who want to deeply study the TW error source including the diurnals, calibration variations, TW triangle behaviors, time link comparisons etc. Those who are not interesting in these details can ignore this section.

### 4.1 Statistics of the triangle closures

The last BIPM calibration was in Aug. 2009. All the UTC calibrations were restored after the satellite change with the method of the GPS bridge. All the non UTC links in the TW network, except for AOS, IPQ and NPL, were recalibrated with the method TCC [1].

Now 13 months after, are the TW calibrations change and how much if yes ? As we know that, after the calibration, the triangle closures should be zero and non zero closures are the errors and they are the “true errors”. Table 4 and the Annex 5 illustrate statistic results. In Table 4, the d is the non zero closure of the triangle composed by the link in the table with the two adjacent UTC links. All the d concerned links with AOS, IPQ, NPL and USNO, for example AOS-NIST, demonstrate the residuals of this computation. All the others, for example, NIST-OP, demonstrate the variations (in addition to the computation residuals) since last calibration (Aug. 2009) of the calibrations of the related three links in the triangle of NIST-OP-PTB. Annex 5 illustrates the histograms of the closures. The data used are between MJD 55430 to 55469.

The TCC uncertainty is 2 ns, cf. [1,2]. We observe that in the Table 4 and Annex 5, most of the d are less than 1 ns. Considering the uA of TW is 0.5 and the TW quality is degraded since the switch to the new satellite, the calibration is very stable. A few (in blue) are bigger than 1 ns, by looking at the plots in Annex 2, 3 and 4, we think that these anomaly closures are from some causes other than the TW measurements or calibrations, e.g. some IT concerned links were obviously biased by an inconstant slop (e.g. the last plot in the first page of Annex 3).

Histogram - Link Triangle Closure of T PTB CH USNO , Mean/RMS/SD=   -4  731  730

PicoSec|No.|     Total In/Out:  456/  416

-2250|  2| ++

-2000|  1| +

-1750|  5| +++++

-1500|  7| +++++++

-1250| 13| +++++++++++++

-1000| 17| +++++++++++++++++

-750| 38| ++++++++++++++++++++++++++++++++++++++

-500| 38| ++++++++++++++++++++++++++++++++++++++

-250| 56| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 56| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 52| ++++++++++++++++++++++++++++++++++++++++++++++++++++

750| 37| +++++++++++++++++++++++++++++++++++++

1000| 18| ++++++++++++++++++

1250| 10| ++++++++++

1500|  7| +++++++

1750|  2| ++

2000|  2| ++

2250|  1| +

### Figure 4.1 the histogram of the triangle closures of PTB, CH and USNO

### Table 4. Verification of the new CALR values with the Triangle closures data of TW1008 and 1009 MJD 55432-55469

### 4.2 Comparison of the TW and GPS PPP UTC time links

The following table gives the differences with the UTC TW links and that of GPS. Annex 3 illustrates the plot of these link comparisons.

The UTC TW links between PTB and CH, IT, OP, SP, USNO and VSL (blue in the following table) are equipment calibrated with uncertainty uB about 1ns. On the other side, the GPS P3 (PPP) were calibrated with uB=5ns. As illustrated in the Table 5 and in the Annex 3, the differences between the links of PTB with IT (PPP), OP, USNO and VSL are less or about 2ns. In particular, as pointed out above, the USNO-PTB TW KU link was aligned to the X link in last November 2009. The difference between the KU and PPP is now only 1.2 ns. This proves the alignment operation in the early of 2010 of the 3ns correction was correct.

Some other links have differences bigger than 5 ns, such as PTB with CH, ROA and SP. The difference between TW and GPS PPP of the baseline of ROA-PTB is 5.1 ns which is rather different than an earlier report [3].

It seems a little strange that for the baseline IT-PTB, the difference between PPP and P3 is 4.7-2.1=2.6 ns while both should have the same calibration ! That was why, I think there is a bias in the related links (see section 4.2 and the last plot of the first page of Annex 3 as well as the plot just after).

### Table 5 Differences between TW and GPS over the UTC baselines

0.2  0.0 AOSPTB.TTTT_.lkT   +-  1.2    363     ! by AOSPTB.TMTA4

6.9  0.0 CHPTB.TTTT_.lkT    +-  0.8    382     ! by CHPTB.TGT34

-4.7  0.0 ITPTB.TTTT_.lkT    +-  1.9    342     ! by ITPTB.TPTA4

-2.1  0.0 ITPTB.TTTT_.lkT    +-  1.9    342     ! by ITPTB.TGT34

2.9  0.0 NISTPTB.TTTT_.lkT  +-  0.4    401     ! by NISTPTB.TGT34

-1.1  0.0 NPLPTB.TTTT_.lkT   +-  0.4    118     ! by NPLPTB.TGT34

-0.6  0.0 OPPTB.TTTT_.lkT    +-  0.7    400     ! by OPPTB.TGT34

5.1  0.0 ROAPTB.TTTT_.lkT   +-  1.2    391     ! by ROAPTB.TGT34

6.2  0.0 SPPTB.TTTT_.lkT    +-  0.8    388     ! by SPPTB.TGT34

1.2  0.0 USNOPTB.TTTT_.lkT  +-  1.3    372     ! by USNOPTB.TGT34

2.6  0.0 VSLPTB.TTTT_.lkT   +-  1.2    372     ! by VSLPTB.TGT34

END ---1009 dLk=Lk1-Lk2=Align Lk1 to Lk2=M1(Lab2-Lab1/Lk1)=CLB(Lab2ofLk1) by F3s34/Z4

### 4.3 Comparison of the TW and GPS PPP Non UTC time links

Similar to the Annex 3, the Annex 4 gives the link comparisons for the Non-UTC baselines. The differences are that in Annex 3, the TW links are of the original calibrations and in Annex 4, the links are of the TTC calibration.

## 5. References:

[1] Jiang, Lewandowski and Piester, BIPM TM 166, Restoration the TWSTFT link calibration using GPS bridging after the satellite switch from IS-3R to T-11N in July/August 2009,Aug 2009

[2] Jiang, Piester and Liang,  Restoring a TWSTFT Calibration with a GPS Bridge - A standard procedure for UTC time transfer,  Proc. EFTF 2010

[3] H. Esteban, J. Palacio, F. J. Galindo, T. Feldmann, A. Bauch, D. Piester, A GPS calibration trip expeerience between ROA and PTB, Proc. EFTF 2009

## Annex 1. How to make the Implementation

### A1.1 ITU data file header entries

Here are listed examples to be implemented into the “TW” file header. Because the UTC-links are bridged links the header entry should be “CAL 125 BRIDGED” (see CH01 – PTB01), the non-UTC links are recalibrated by the TCC and thus the header entry should be “TRIANGLE CLOSURE”(see CH01 - IT01).

Example 1.

European “TW”-files:

* LINK   12 SAT: TELSTAR 11N         NLO: E 322 30 00.000  XPNDR: 999999999 ns

*           SAT-NTX: 11004.7200 MHz  SAT-NRX: 14304.7200 MHz  BW:   2.5 MHz

* LINK   13 SAT: TELSTAR 11N         NLO: E 322 30 00.000  XPNDR: 999999999 ns

*           SAT-NTX: 11485.5500 MHz  SAT-NRX: 14026.5500 MHz  BW:   2.5 MHz

US “TW”-files:

* LINK   13 SAT: TELSTAR 11N         NLO: E 322 30 00.000  XPNDR: 999999999 ns

*           SAT-NTX: 11726.5500 MHz  SAT-NRX: 14285.5500 MHz  BW:   2.5 MHz

Examples for the additional header entries:

CH01 – PTB01:

* CAL   139 TYPE: CAL 125 BRIDGED    MJD: 55043  EST. UNCERT.:    1.000 ns

USNO01 – PTB01:

* CAL   140 TYPE: CAL 114 BRIDGED    MJD: 55043  EST. UNCERT.:    1.100 ns

NIST01 – PTB01:

* CAL   141 TYPE: CAL 113 BRIDGED    MJD: 55043  EST. UNCERT.:    5.000 ns

IT01 – PTB01:

* CAL   142 TYPE: CAL 113 BRIDGED    MJD: 55043  EST. UNCERT.:    1.200 ns

OP01 – PTB01:

* CAL   143 TYPE: CAL 134 BRIDGED    MJD: 55043  EST. UNCERT.:    1.100 ns

ROA01 – PTB01:

* CAL   144 TYPE: CAL 118 BRIDGED    MJD: 55043  EST. UNCERT.:    5.000 ns

SP01 – PTB01:

* CAL   145 TYPE: CAL 119 BRIDGED    MJD: 55043  EST. UNCERT.:    1.100 ns

VSL01 – PTB01:

* CAL   146 TYPE: CAL 120 BRIDGED    MJD: 55043  EST. UNCERT.:    1.200 ns

CH01 – IT01:

* CAL   147 TYPE: TRIANGLE CLOSURE   MJD: 55043  EST. UNCERT.:    2.000 ns

### Example 2. the header of IPQ)

1  * FORMAT    01 
2  * LAB       IPQ 
3  * REV DATE  2010-11-02 
4  * ES IPQ01  LA: N  38 40 16.023      LO: W   9 11 10.033   HT: +0151.20 m 
5  * REF-FRAME ITRF 2005 
6  * LINK   12 SAT: TELSTAR 11N         NLO: E 322 30 00.000  XPNDR: 999999999 ns 
7  *           SAT-NTX: 11004.7200 MHz  SAT-NRX: 14304.7200 MHz  BW:   2.5 MHz 
8  * LINK   13 SAT: TELSTAR 11N         NLO: E 322 30 00.000  XPNDR: 999999999 ns 
9  *           SAT-NTX: 11485.5500 MHz  SAT-NRX: 14026.5500 MHz  BW:   2.5 MHz 
18 * CAL   172 TYPE: GPS                MJD: 55502  EST. UNCERT.:      7.0 ns 
10 * CAL   177 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
11 * CAL   184 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
12 * CAL   187 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
13 * CAL   190 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
14 * CAL   191 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
15 * CAL   192 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
16 * CAL   194 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
17 * CAL   195 TYPE: TRIANGLE CLOSURE   MJD: 55502  EST. UNCERT.:      7.0 ns 
20 * LOC-MON   NO 
21 * MODEM     SATRE 318 
22 * COMMENTS:

### A1.2 Examples for ITU-format data recordings

Examples for the data lines are listed in the following. Changes to be applied are in bold and blue numbers.

* EARTH-STAT  LI  MJD  STTIME NTL        TW        DRMS SMP ATL     REFDELAY     RSIG  CI S    CALR     ESDVAR   ESIG TMP HUM PRES

* LOC    REM           hhmmss  s         s          ns       s         s          ns            ns        ns      ns degC  %  mbar

TWIPQ55.437

Currently :

IPQ01  PTB01 12 55437 004000 119  0.260344277075 0.114 121 120  0.000000793149 0.090 999 9 999999999 999999999 99999 999 999 9999

Replaced :

IPQ01  PTB01 12 55437 004000 119  0.260344277075 0.114 121 120  0.000000793149 0.090 172 1  -188.600 999999999 99999 999 999 9999

## Annex 2. Link comparisons on the baseline USNO-PTB

## Annex 3. Link comparisons on the UTC baselines

## Annex 4. Link comparisons on the calibrated Non UTC baselines

## Annex 5. Triangle closures of the BIPM 2009 and 2010 calibrated time links

Histogram - Link Triangle Closure of T PTB AOS CH  , Mean/RMS/SD=    5  726  726

PicoSec|No.|     Total In/Out:  456/  388

-2000|  1| +

-1750|  1| +

-1500|  8| ++++++++

-1250|  8| ++++++++

-1000| 21| +++++++++++++++++++++

-750| 35| +++++++++++++++++++++++++++++++++++

-500| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

0| 46| ++++++++++++++++++++++++++++++++++++++++++++++

250| 52| ++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 37| +++++++++++++++++++++++++++++++++++++

750| 26| ++++++++++++++++++++++++++

1000| 17| +++++++++++++++++

1250| 18| ++++++++++++++++++

1500|  6| ++++++

1750|  6| ++++++

2000|  1| +

2250|  1| +

Histogram - Link Triangle Closure of T PTB AOS IT  , Mean/RMS/SD=  -12 1568 1568

PicoSec|No.|     Total In/Out:  456/  372

-3250|  3| +++

-3000|  6| ++++++

-2750|  4| ++++

-2500|  8| ++++++++

-2250|  7| +++++++

-2000| 18| ++++++++++++++++++

-1750| 22| ++++++++++++++++++++++

-1500| 26| ++++++++++++++++++++++++++

-1250| 27| +++++++++++++++++++++++++++

-1000| 14| ++++++++++++++

-750| 15| +++++++++++++++

-500| 10| ++++++++++

-250| 20| ++++++++++++++++++++

0| 20| ++++++++++++++++++++

250| 15| +++++++++++++++

500| 12| ++++++++++++

750| 11| +++++++++++

1000| 17| +++++++++++++++++

1250| 21| +++++++++++++++++++++

1500| 21| +++++++++++++++++++++

1750| 22| ++++++++++++++++++++++

2000| 22| ++++++++++++++++++++++

2250| 15| +++++++++++++++

2500|  8| ++++++++

2750|  4| ++++

3000|  2| ++

3250|  2| ++

Histogram - Link Triangle Closure of T PTB AOS IPQ , Mean/RMS/SD=   -9  795  795

PicoSec|No.|     Total In/Out:  456/  381

-3250|  1| +

-3000|  0|

-2750|  0|

-2500|  0|

-2250|  0|

-2000|  0|

-1750|  2| ++

-1500| 10| ++++++++++

-1250| 13| +++++++++++++

-1000| 26| ++++++++++++++++++++++++++

-750| 39| +++++++++++++++++++++++++++++++++++++++

-500| 43| +++++++++++++++++++++++++++++++++++++++++++

-250| 36| ++++++++++++++++++++++++++++++++++++

0| 42| ++++++++++++++++++++++++++++++++++++++++++

250| 39| +++++++++++++++++++++++++++++++++++++++

500| 50| ++++++++++++++++++++++++++++++++++++++++++++++++++

750| 29| +++++++++++++++++++++++++++++

1000| 22| ++++++++++++++++++++++

1250| 11| +++++++++++

1500|  9| +++++++++

1750|  5| +++++

2000|  4| ++++

Histogram - Link Triangle Closure of T PTB AOS NIST, Mean/RMS/SD=   -1  590  590

PicoSec|No.|     Total In/Out:  456/  390

-2000|  1| +

-1750|  2| ++

-1500|  1| +

-1250|  6| ++++++

-1000|  7| +++++++

-750| 24| ++++++++++++++++++++++++

-500| 65| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 59| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 75| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 59| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 35| +++++++++++++++++++++++++++++++++++

750| 27| +++++++++++++++++++++++++++

1000| 11| +++++++++++

1250| 12| ++++++++++++

1500|  4| ++++

1750|  1| +

2000|  1| +

Histogram - Link Triangle Closure of T PTB AOS OP  , Mean/RMS/SD=    6 1120 1120

PicoSec|No.|     Total In/Out:  456/  390

-3250|  1| +

-3000|  0|

-2750|  0|

-2500|  1| +

-2250|  2| ++

-2000|  3| +++

-1750| 12| ++++++++++++

-1500| 18| ++++++++++++++++++

-1250| 30| ++++++++++++++++++++++++++++++

-1000| 29| +++++++++++++++++++++++++++++

-750| 38| ++++++++++++++++++++++++++++++++++++++

-500| 29| +++++++++++++++++++++++++++++

-250| 32| ++++++++++++++++++++++++++++++++

0| 21| +++++++++++++++++++++

250| 27| +++++++++++++++++++++++++++

500| 25| +++++++++++++++++++++++++

750| 25| +++++++++++++++++++++++++

1000| 28| ++++++++++++++++++++++++++++

1250| 24| ++++++++++++++++++++++++

1500| 13| +++++++++++++

1750|  6| ++++++

2000| 12| ++++++++++++

2250|  7| +++++++

2500|  3| +++

2750|  3| +++

3000|  1| +

Histogram - Link Triangle Closure of T PTB AOS ROA , Mean/RMS/SD=    2  723  723

PicoSec|No.|     Total In/Out:  456/  387

-2000|  1| +

-1750|  3| +++

-1500|  6| ++++++

-1250| 14| ++++++++++++++

-1000| 15| +++++++++++++++

-750| 27| +++++++++++++++++++++++++++

-500| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

-250| 53| +++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 68| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 36| ++++++++++++++++++++++++++++++++++++

500| 46| ++++++++++++++++++++++++++++++++++++++++++++++

750| 31| +++++++++++++++++++++++++++++++

1000| 15| +++++++++++++++

1250|  9| +++++++++

1500|  6| ++++++

1750|  6| ++++++

2000|  0|

2250|  1| +

2500|  0|

2750|  0|

3000|  0|

3250|  1| +

Histogram - Link Triangle Closure of T PTB AOS SP  , Mean/RMS/SD=   15 1199 1199

PicoSec|No.|     Total In/Out:  456/  375

-3250|  1| +

-3000|  3| +++

-2750|  3| +++

-2500|  4| ++++

-2250|  4| ++++

-2000|  5| +++++

-1750| 14| ++++++++++++++

-1500| 15| +++++++++++++++

-1250| 24| ++++++++++++++++++++++++

-1000| 15| +++++++++++++++

-750| 19| +++++++++++++++++++

-500| 31| +++++++++++++++++++++++++++++++

-250| 33| +++++++++++++++++++++++++++++++++

0| 24| ++++++++++++++++++++++++

250| 32| ++++++++++++++++++++++++++++++++

500| 32| ++++++++++++++++++++++++++++++++

750| 30| ++++++++++++++++++++++++++++++

1000| 23| +++++++++++++++++++++++

1250| 12| ++++++++++++

1500| 20| ++++++++++++++++++++

1750| 12| ++++++++++++

2000|  4| ++++

2250|  3| +++

2500|  7| +++++++

2750|  2| ++

3000|  2| ++

3250|  1| +

Histogram - Link Triangle Closure of T PTB AOS USNO, Mean/RMS/SD=    5  762  762

PicoSec|No.|     Total In/Out:  456/  379

-2000|  3| +++

-1750| 10| ++++++++++

-1500|  4| ++++

-1250|  9| +++++++++

-1000| 28| ++++++++++++++++++++++++++++

-750| 26| ++++++++++++++++++++++++++

-500| 36| ++++++++++++++++++++++++++++++++++++

-250| 33| +++++++++++++++++++++++++++++++++

0| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

250| 57| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 48| ++++++++++++++++++++++++++++++++++++++++++++++++

750| 29| +++++++++++++++++++++++++++++

1000| 19| +++++++++++++++++++

1250| 13| +++++++++++++

1500|  9| +++++++++

1750|  4| ++++

Histogram - Link Triangle Closure of T PTB AOS VSL , Mean/RMS/SD=  -13 1324 1324

PicoSec|No.|     Total In/Out:  456/  328

-3750|  1| +

-3500|  0|

-3250|  0|

-3000|  1| +

-2750|  4| ++++

-2500|  4| ++++

-2250|  5| +++++

-2000| 10| ++++++++++

-1750| 12| ++++++++++++

-1500| 13| +++++++++++++

-1250| 23| +++++++++++++++++++++++

-1000| 20| ++++++++++++++++++++

-750| 24| ++++++++++++++++++++++++

-500| 18| ++++++++++++++++++

-250| 27| +++++++++++++++++++++++++++

0| 22| ++++++++++++++++++++++

250| 22| ++++++++++++++++++++++

500| 13| +++++++++++++

750| 19| +++++++++++++++++++

1000| 18| ++++++++++++++++++

1250| 13| +++++++++++++

1500| 13| +++++++++++++

1750| 17| +++++++++++++++++

2000|  8| ++++++++

2250| 10| ++++++++++

2500|  6| ++++++

2750|  4| ++++

3000|  1| +

Histogram - Link Triangle Closure of T PTB CH IT   , Mean/RMS/SD= 2186 2668 1530

PicoSec|No.|     Total In/Out:  456/  399

-2500|  1| +

-2250|  0|

-2000|  0|

-1750|  0|

-1500|  0|

-1250|  3| +++

-1000|  8| ++++++++

-750|  5| +++++

-500|  4| ++++

-250|  8| ++++++++

0| 13| +++++++++++++

250| 10| ++++++++++

500| 27| +++++++++++++++++++++++++++

750| 21| +++++++++++++++++++++

1000| 16| ++++++++++++++++

1250| 14| ++++++++++++++

1500| 17| +++++++++++++++++

1750| 15| +++++++++++++++

2000| 18| ++++++++++++++++++

2250| 22| ++++++++++++++++++++++

2500| 22| ++++++++++++++++++++++

2750| 22| ++++++++++++++++++++++

3000| 24| ++++++++++++++++++++++++

3250| 21| +++++++++++++++++++++

3500| 25| +++++++++++++++++++++++++

3750| 31| +++++++++++++++++++++++++++++++

4000| 21| +++++++++++++++++++++

4250| 17| +++++++++++++++++

4500| 12| ++++++++++++

4750|  2| ++

Histogram - Link Triangle Closure of T PTB CH IPQ  , Mean/RMS/SD=  -15 1056 1056

PicoSec|No.|     Total In/Out:  456/  419

-3500|  1| +

-3250|  0|

-3000|  0|

-2750|  0|

-2500|  1| +

-2250|  4| ++++

-2000|  8| ++++++++

-1750| 13| +++++++++++++

-1500| 14| ++++++++++++++

-1250| 27| +++++++++++++++++++++++++++

-1000| 29| +++++++++++++++++++++++++++++

-750| 32| ++++++++++++++++++++++++++++++++

-500| 29| +++++++++++++++++++++++++++++

-250| 35| +++++++++++++++++++++++++++++++++++

0| 32| ++++++++++++++++++++++++++++++++

250| 36| ++++++++++++++++++++++++++++++++++++

500| 37| +++++++++++++++++++++++++++++++++++++

750| 30| ++++++++++++++++++++++++++++++

1000| 25| +++++++++++++++++++++++++

1250| 26| ++++++++++++++++++++++++++

1500| 14| ++++++++++++++

1750| 13| +++++++++++++

2000|  6| ++++++

2250|  4| ++++

2500|  1| +

2750|  1| +

3000|  1| +

Histogram - Link Triangle Closure of T PTB CH NIST , Mean/RMS/SD= -449  818  684

PicoSec|No.|     Total In/Out:  456/  431

-3250|  1| +

-3000|  0|

-2750|  0|

-2500|  0|

-2250|  3| +++

-2000|  7| +++++++

-1750| 13| +++++++++++++

-1500| 16| ++++++++++++++++

-1250| 33| +++++++++++++++++++++++++++++++++

-1000| 45| +++++++++++++++++++++++++++++++++++++++++++++

-750| 43| +++++++++++++++++++++++++++++++++++++++++++

-500| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 60| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 62| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 53| +++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 26| ++++++++++++++++++++++++++

750| 10| ++++++++++

1000|  4| ++++

1250|  1| +

Histogram - Link Triangle Closure of T PTB CH NPL  , Mean/RMS/SD=  -27  951  950

PicoSec|No.|     Total In/Out:  456/  118

-2750|  1| +

-2500|  0|

-2250|  1| +

-2000|  1| +

-1750|  4| ++++

-1500|  1| +

-1250|  2| ++

-1000|  9| +++++++++

-750|  8| ++++++++

-500| 16| ++++++++++++++++

-250| 16| ++++++++++++++++

0|  9| +++++++++

250| 14| ++++++++++++++

500|  6| ++++++

750|  9| +++++++++

1000|  5| +++++

1250|  8| ++++++++

1500|  4| ++++

1750|  2| ++

2000|  0|

2250|  1| +

2500|  0|

2750|  1| +

Histogram - Link Triangle Closure of T PTB CH OP   , Mean/RMS/SD=  150  829  815

PicoSec|No.|     Total In/Out:  456/  430

-2250|  1| +

-2000|  2| ++

-1750|  3| +++

-1500|  6| ++++++

-1250| 17| +++++++++++++++++

-1000| 18| ++++++++++++++++++

-750| 34| ++++++++++++++++++++++++++++++++++

-500| 36| ++++++++++++++++++++++++++++++++++++

-250| 42| ++++++++++++++++++++++++++++++++++++++++++

0| 47| +++++++++++++++++++++++++++++++++++++++++++++++

250| 41| +++++++++++++++++++++++++++++++++++++++++

500| 52| ++++++++++++++++++++++++++++++++++++++++++++++++++++

750| 41| +++++++++++++++++++++++++++++++++++++++++

1000| 43| +++++++++++++++++++++++++++++++++++++++++++

1250| 23| +++++++++++++++++++++++

1500| 12| ++++++++++++

1750|  9| +++++++++

2000|  1| +

2250|  2| ++

Histogram - Link Triangle Closure of T PTB CH ROA  , Mean/RMS/SD= -685 1008  739

PicoSec|No.|     Total In/Out:  456/  427

-2750|  1| +

-2500|  4| ++++

-2250|  8| ++++++++

-2000| 12| ++++++++++++

-1750| 15| +++++++++++++++

-1500| 29| +++++++++++++++++++++++++++++

-1250| 43| +++++++++++++++++++++++++++++++++++++++++++

-1000| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

-750| 74| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-500| 58| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 47| +++++++++++++++++++++++++++++++++++++++++++++++

0| 35| +++++++++++++++++++++++++++++++++++

250| 16| ++++++++++++++++

500| 12| ++++++++++++

750| 13| +++++++++++++

1000|  6| ++++++

1250|  3| +++

1500|  2| ++

Histogram - Link Triangle Closure of T PTB CH SP   , Mean/RMS/SD=  285 1176 1140

PicoSec|No.|     Total In/Out:  456/  371

-3000|  1| +

-2750|  1| +

-2500|  2| ++

-2250|  2| ++

-2000|  4| ++++

-1750|  3| +++

-1500| 13| +++++++++++++

-1250| 14| ++++++++++++++

-1000| 23| +++++++++++++++++++++++

-750| 22| ++++++++++++++++++++++

-500| 21| +++++++++++++++++++++

-250| 21| +++++++++++++++++++++

0| 28| ++++++++++++++++++++++++++++

250| 27| +++++++++++++++++++++++++++

500| 35| +++++++++++++++++++++++++++++++++++

750| 40| ++++++++++++++++++++++++++++++++++++++++

1000| 27| +++++++++++++++++++++++++++

1250| 20| ++++++++++++++++++++

1500| 24| ++++++++++++++++++++++++

1750| 19| +++++++++++++++++++

2000| 13| +++++++++++++

2250|  4| ++++

2500|  4| ++++

2750|  1| +

3000|  2| ++

Histogram - Link Triangle Closure of T PTB CH USNO , Mean/RMS/SD=   -4  731  730

PicoSec|No.|     Total In/Out:  456/  416

-2250|  2| ++

-2000|  1| +

-1750|  5| +++++

-1500|  7| +++++++

-1250| 13| +++++++++++++

-1000| 17| +++++++++++++++++

-750| 38| ++++++++++++++++++++++++++++++++++++++

-500| 38| ++++++++++++++++++++++++++++++++++++++

-250| 56| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 56| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 52| ++++++++++++++++++++++++++++++++++++++++++++++++++++

750| 37| +++++++++++++++++++++++++++++++++++++

1000| 18| ++++++++++++++++++

1250| 10| ++++++++++

1500|  7| +++++++

1750|  2| ++

2000|  2| ++

2250|  1| +

Histogram - Link Triangle Closure of T PTB CH VSL  , Mean/RMS/SD= -658 1269 1085

PicoSec|No.|     Total In/Out:  456/   98

-3500|  1| +

-3250|  0|

-3000|  0|

-2750|  2| ++

-2500|  1| +

-2250|  1| +

-2000|  6| ++++++

-1750|  5| +++++

-1500|  6| ++++++

-1250| 10| ++++++++++

-1000| 12| ++++++++++++

-750| 13| +++++++++++++

-500|  7| +++++++

-250|  7| +++++++

0|  8| ++++++++

250|  3| +++

500|  2| ++

750|  3| +++

1000|  4| ++++

1250|  2| ++

1500|  3| +++

1750|  0|

2000|  1| +

2250|  1| +

Histogram - Link Triangle Closure of T PTB IT IPQ  , Mean/RMS/SD=    9 1526 1526

PicoSec|No.|     Total In/Out:  456/  318

-4750|  1| +

-4500|  0|

-4250|  0|

-4000|  0|

-3750|  1| +

-3500|  2| ++

-3250|  0|

-3000|  3| +++

-2750|  0|

-2500|  3| +++

-2250|  7| +++++++

-2000| 10| ++++++++++

-1750| 17| +++++++++++++++++

-1500| 17| +++++++++++++++++

-1250| 14| ++++++++++++++

-1000| 22| ++++++++++++++++++++++

-750| 22| ++++++++++++++++++++++

-500| 27| +++++++++++++++++++++++++++

-250| 12| ++++++++++++

0| 14| ++++++++++++++

250| 17| +++++++++++++++++

500| 19| +++++++++++++++++++

750| 21| +++++++++++++++++++++

1000|  9| +++++++++

1250| 10| ++++++++++

1500| 19| +++++++++++++++++++

1750| 12| ++++++++++++

2000| 11| +++++++++++

2250|  8| ++++++++

2500|  3| +++

2750|  4| ++++

3000|  8| ++++++++

3250|  2| ++

3500|  1| +

3750|  0|

4000|  1| +

4250|  0|

4500|  1| +

Histogram - Link Triangle Closure of T PTB IT NIST , Mean/RMS/SD=-3131 3318 1098

PicoSec|No.|     Total In/Out:  456/  395

-5000|  7| +++++++

-4750| 18| ++++++++++++++++++

-4500| 22| ++++++++++++++++++++++

-4250| 34| ++++++++++++++++++++++++++++++++++

-4000| 33| +++++++++++++++++++++++++++++++++

-3750| 23| +++++++++++++++++++++++

-3500| 31| +++++++++++++++++++++++++++++++

-3250| 28| ++++++++++++++++++++++++++++

-3000| 24| ++++++++++++++++++++++++

-2750| 27| +++++++++++++++++++++++++++

-2500| 34| ++++++++++++++++++++++++++++++++++

-2250| 27| +++++++++++++++++++++++++++

-2000| 28| ++++++++++++++++++++++++++++

-1750| 26| ++++++++++++++++++++++++++

-1500| 15| +++++++++++++++

-1250| 10| ++++++++++

-1000|  3| +++

-750|  3| +++

-500|  0|

-250|  1| +

0|  1| +

Histogram - Link Triangle Closure of T PTB IT NPL  , Mean/RMS/SD=  -11  573  573

PicoSec|No.|     Total In/Out:  456/  230

-1250|  3| +++

-1000|  4| ++++

-750| 24| ++++++++++++++++++++++++

-500| 33| +++++++++++++++++++++++++++++++++

-250| 41| +++++++++++++++++++++++++++++++++++++++++

0| 44| ++++++++++++++++++++++++++++++++++++++++++++

250| 31| +++++++++++++++++++++++++++++++

500| 20| ++++++++++++++++++++

750| 11| +++++++++++

1000| 10| ++++++++++

1250|  6| ++++++

1500|  1| +

1750|  0|

2000|  1| +

2250|  1| +

Histogram - Link Triangle Closure of T PTB IT OP   , Mean/RMS/SD=-1120 1437  900

PicoSec|No.|     Total In/Out:  456/  403

-3000|  1| +

-2750|  5| +++++

-2500|  9| +++++++++

-2250| 35| +++++++++++++++++++++++++++++++++++

-2000| 33| +++++++++++++++++++++++++++++++++

-1750| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

-1500| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

-1250| 43| +++++++++++++++++++++++++++++++++++++++++++

-1000| 33| +++++++++++++++++++++++++++++++++

-750| 26| ++++++++++++++++++++++++++

-500| 16| ++++++++++++++++

-250| 18| ++++++++++++++++++

0| 26| ++++++++++++++++++++++++++

250| 11| +++++++++++

500| 18| ++++++++++++++++++

750| 13| +++++++++++++

1000|  4| ++++

1250|  2| ++

Histogram - Link Triangle Closure of T PTB IT ROA  , Mean/RMS/SD= -478  563  296

PicoSec|No.|     Total In/Out:  456/  400

-1500|  3| +++

-1250|  1| +

-1000| 27| +++++++++++++++++++++++++++

-750| 89| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-500|144| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 91| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 37| +++++++++++++++++++++++++++++++++++++

250|  6| ++++++

500|  0|

750|  0|

1000|  2| ++

Histogram - Link Triangle Closure of T PTB IT SP   , Mean/RMS/SD= -740 1126  849

PicoSec|No.|     Total In/Out:  456/  397

-2500|  2| ++

-2250|  7| +++++++

-2000| 13| +++++++++++++

-1750| 25| +++++++++++++++++++++++++

-1500| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

-1250| 61| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-1000| 61| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-750| 30| ++++++++++++++++++++++++++++++

-500| 23| +++++++++++++++++++++++

-250| 17| +++++++++++++++++

0| 32| ++++++++++++++++++++++++++++++++

250| 28| ++++++++++++++++++++++++++++

500| 14| ++++++++++++++

750| 19| +++++++++++++++++++

1000| 11| +++++++++++

1250|  4| ++++

1500|  0|

1750|  0|

2000|  1| +

Histogram - Link Triangle Closure of T PTB IT USNO , Mean/RMS/SD=   13 1732 1732

PicoSec|No.|     Total In/Out:  456/  391

-4750|  1| +

-4500|  1| +

-4250|  0|

-4000|  0|

-3750|  1| +

-3500|  2| ++

-3250|  3| +++

-3000|  5| +++++

-2750|  6| ++++++

-2500|  5| +++++

-2250| 12| ++++++++++++

-2000| 15| +++++++++++++++

-1750| 15| +++++++++++++++

-1500| 17| +++++++++++++++++

-1250| 27| +++++++++++++++++++++++++++

-1000| 23| +++++++++++++++++++++++

-750| 30| ++++++++++++++++++++++++++++++

-500| 21| +++++++++++++++++++++

-250| 29| +++++++++++++++++++++++++++++

0| 11| +++++++++++

250| 13| +++++++++++++

500| 18| ++++++++++++++++++

750| 10| ++++++++++

1000| 12| ++++++++++++

1250| 11| +++++++++++

1500| 19| +++++++++++++++++++

1750| 14| ++++++++++++++

2000| 13| +++++++++++++

2250| 10| ++++++++++

2500| 15| +++++++++++++++

2750| 13| +++++++++++++

3000|  7| +++++++

3250|  3| +++

3500|  2| ++

3750|  5| +++++

4000|  1| +

4250|  1| +

Histogram - Link Triangle Closure of T PTB IT VSL  , Mean/RMS/SD= -544  616  288

PicoSec|No.|     Total In/Out:  456/   10

-1000|  2| ++

-750|  1| +

-500|  3| +++

-250|  4| ++++

Histogram - Link Triangle Closure of T PTB IPQ NIST, Mean/RMS/SD=    5  726  726

PicoSec|No.|     Total In/Out:  456/  419

-1750|  4| ++++

-1500| 10| ++++++++++

-1250| 16| ++++++++++++++++

-1000| 23| +++++++++++++++++++++++

-750| 24| ++++++++++++++++++++++++

-500| 45| +++++++++++++++++++++++++++++++++++++++++++++

-250| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

0| 57| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 67| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 45| +++++++++++++++++++++++++++++++++++++++++++++

750| 30| ++++++++++++++++++++++++++++++

1000| 25| +++++++++++++++++++++++++

1250|  9| +++++++++

1500|  8| ++++++++

1750|  4| ++++

2000|  0|

2250|  0|

2500|  0|

2750|  0|

3000|  1| +

Histogram - Link Triangle Closure of T PTB IPQ OP  , Mean/RMS/SD=   -2 1194 1194

PicoSec|No.|     Total In/Out:  456/  424

-4250|  1| +

-4000|  0|

-3750|  0|

-3500|  0|

-3250|  1| +

-3000|  0|

-2750|  0|

-2500|  3| +++

-2250|  6| ++++++

-2000|  3| +++

-1750| 11| +++++++++++

-1500| 19| +++++++++++++++++++

-1250| 20| ++++++++++++++++++++

-1000| 41| +++++++++++++++++++++++++++++++++++++++++

-750| 32| ++++++++++++++++++++++++++++++++

-500| 41| +++++++++++++++++++++++++++++++++++++++++

-250| 34| ++++++++++++++++++++++++++++++++++

0| 41| +++++++++++++++++++++++++++++++++++++++++

250| 30| ++++++++++++++++++++++++++++++

500| 27| +++++++++++++++++++++++++++

750| 25| +++++++++++++++++++++++++

1000| 13| +++++++++++++

1250| 13| +++++++++++++

1500| 17| +++++++++++++++++

1750|  9| +++++++++

2000| 10| ++++++++++

2250| 11| +++++++++++

2500|  6| ++++++

2750|  6| ++++++

3000|  0|

3250|  3| +++

3500|  1| +

Histogram - Link Triangle Closure of T PTB IPQ ROA , Mean/RMS/SD=   17  853  853

PicoSec|No.|     Total In/Out:  456/  420

-2750|  1| +

-2500|  0|

-2250|  1| +

-2000|  4| ++++

-1750|  3| +++

-1500|  7| +++++++

-1250| 16| ++++++++++++++++

-1000| 18| ++++++++++++++++++

-750| 41| +++++++++++++++++++++++++++++++++++++++++

-500| 36| ++++++++++++++++++++++++++++++++++++

-250| 70| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 40| ++++++++++++++++++++++++++++++++++++++++

500| 38| ++++++++++++++++++++++++++++++++++++++

750| 28| ++++++++++++++++++++++++++++

1000| 21| +++++++++++++++++++++

1250| 14| ++++++++++++++

1500|  7| +++++++

1750| 14| ++++++++++++++

2000|  4| ++++

2250|  1| +

2500|  1| +

2750|  0|

3000|  0|

3250|  0|

3500|  0|

3750|  0|

4000|  0|

4250|  0|

4500|  1| +

Histogram - Link Triangle Closure of T PTB IPQ SP  , Mean/RMS/SD=   13 3763 3763

PicoSec|No.|     Total In/Out:  456/  225

-5000|  3| +++

-4750|  3| +++

-4500|  2| ++

-4250|  2| ++

-4000|  2| ++

-3750|  3| +++

-3500|  5| +++++

-3250|  4| ++++

-3000|  6| ++++++

-2750|  5| +++++

-2500|  5| +++++

-2250| 11| +++++++++++

-2000|  7| +++++++

-1750| 10| ++++++++++

-1500| 12| ++++++++++++

-1250| 10| ++++++++++

-1000|  7| +++++++

-750| 10| ++++++++++

-500|  5| +++++

-250|  9| +++++++++

0| 10| ++++++++++

250|  9| +++++++++

500|  7| +++++++

750|  4| ++++

1000|  3| +++

1250|  6| ++++++

1500|  6| ++++++

1750| 10| ++++++++++

2000|  7| +++++++

2250|  5| +++++

2500|  2| ++

2750|  7| +++++++

3000|  2| ++

3250|  5| +++++

3500|  4| ++++

3750|  4| ++++

4000|  2| ++

4250|  5| +++++

4500|  3| +++

4750|  3| +++

Histogram - Link Triangle Closure of T PTB IPQ USNO, Mean/RMS/SD=    9  953  953

PicoSec|No.|     Total In/Out:  456/  391

-2250|  2| ++

-2000|  2| ++

-1750|  5| +++++

-1500| 11| +++++++++++

-1250| 11| +++++++++++

-1000| 24| ++++++++++++++++++++++++

-750| 39| +++++++++++++++++++++++++++++++++++++++

-500| 38| ++++++++++++++++++++++++++++++++++++++

-250| 46| ++++++++++++++++++++++++++++++++++++++++++++++

0| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

250| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 35| +++++++++++++++++++++++++++++++++++

750| 27| +++++++++++++++++++++++++++

1000| 14| ++++++++++++++

1250| 11| +++++++++++

1500|  8| ++++++++

1750|  4| ++++

2000|  3| +++

2250|  3| +++

2500|  0|

2750|  2| ++

3000|  1| +

3250|  0|

3500|  0|

3750|  0|

4000|  0|

4250|  0|

4500|  1| +

Histogram - Link Triangle Closure of T PTB IPQ VSL , Mean/RMS/SD= -137 9107 9106

PicoSec|No.|     Total In/Out:  456/   64

-5000|  3| +++

-4750|  0|

-4500|  1| +

-4250|  1| +

-4000|  1| +

-3750|  0|

-3500|  3| +++

-3250|  1| +

-3000|  4| ++++

-2750|  0|

-2500|  6| ++++++

-2250|  1| +

-2000|  0|

-1750|  1| +

-1500|  2| ++

-1250|  0|

-1000|  1| +

-750|  1| +

-500|  3| +++

-250|  1| +

0|  1| +

250|  0|

500|  2| ++

750|  4| ++++

1000|  2| ++

1250|  3| +++

1500|  0|

1750|  1| +

2000|  2| ++

2250|  1| +

2500|  1| +

2750|  2| ++

3000|  3| +++

3250|  0|

3500|  2| ++

3750|  1| +

4000|  5| +++++

4250|  0|

4500|  2| ++

4750|  2| ++

Histogram - Link Triangle Closure of T PTB NIST OP , Mean/RMS/SD=  424  813  693

PicoSec|No.|     Total In/Out:  456/  421

-1000|  2| ++

-750| 18| ++++++++++++++++++

-500| 42| ++++++++++++++++++++++++++++++++++++++++++

-250| 44| ++++++++++++++++++++++++++++++++++++++++++++

0| 50| ++++++++++++++++++++++++++++++++++++++++++++++++++

250| 48| ++++++++++++++++++++++++++++++++++++++++++++++++

500| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

750| 52| ++++++++++++++++++++++++++++++++++++++++++++++++++++

1000| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

1250| 20| ++++++++++++++++++++

1500| 16| ++++++++++++++++

1750| 19| +++++++++++++++++++

2000|  4| ++++

2250|  2| ++

Histogram - Link Triangle Closure of T PTB NIST ROA, Mean/RMS/SD=    0  661  661

PicoSec|No.|     Total In/Out:  456/  427

-1750|  1| +

-1500|  6| ++++++

-1250| 13| +++++++++++++

-1000| 20| ++++++++++++++++++++

-750| 33| +++++++++++++++++++++++++++++++++

-500| 58| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 53| +++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 57| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 66| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 41| +++++++++++++++++++++++++++++++++++++++++

750| 34| ++++++++++++++++++++++++++++++++++

1000| 26| ++++++++++++++++++++++++++

1250| 12| ++++++++++++

1500|  5| +++++

1750|  1| +

2000|  1| +

Histogram - Link Triangle Closure of T PTB NIST SP , Mean/RMS/SD= 1733 1915  817

PicoSec|No.|     Total In/Out:  456/  424

-1250|  1| +

-1000|  0|

-750|  1| +

-500|  2| ++

-250|  1| +

0| 12| ++++++++++++

250| 16| ++++++++++++++++

500| 12| ++++++++++++

750| 18| ++++++++++++++++++

1000| 25| +++++++++++++++++++++++++

1250| 41| +++++++++++++++++++++++++++++++++++++++++

1500| 49| +++++++++++++++++++++++++++++++++++++++++++++++++

1750| 57| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++

2000| 35| +++++++++++++++++++++++++++++++++++

2250| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

2500| 50| ++++++++++++++++++++++++++++++++++++++++++++++++++

2750| 24| ++++++++++++++++++++++++

3000| 22| ++++++++++++++++++++++

3250|  6| ++++++

3500|  1| +

Histogram - Link Triangle Closure of T PTB NIST VSL, Mean/RMS/SD=  316 1243 1202

PicoSec|No.|     Total In/Out:  456/  369

-2750|  4| ++++

-2500|  4| ++++

-2250|  3| +++

-2000|  5| +++++

-1750|  8| ++++++++

-1500| 14| ++++++++++++++

-1250|  7| +++++++

-1000| 18| ++++++++++++++++++

-750| 18| ++++++++++++++++++

-500| 14| ++++++++++++++

-250| 30| ++++++++++++++++++++++++++++++

0| 22| ++++++++++++++++++++++

250| 33| +++++++++++++++++++++++++++++++++

500| 24| ++++++++++++++++++++++++

750| 43| +++++++++++++++++++++++++++++++++++++++++++

1000| 35| +++++++++++++++++++++++++++++++++++

1250| 29| +++++++++++++++++++++++++++++

1500| 15| +++++++++++++++

1750| 12| ++++++++++++

2000|  9| +++++++++

2250|  6| ++++++

2500|  4| ++++

2750|  4| ++++

3000|  4| ++++

3250|  2| ++

3500|  1| +

3750|  1| +

Histogram - Link Triangle Closure of T PTB NPL OP  , Mean/RMS/SD=   -5  471  471

PicoSec|No.|     Total In/Out:  456/   71

-2000|  1| +

-1750|  0|

-1500|  1| +

-1250|  0|

-1000|  0|

-750|  0|

-500|  4| ++++

-250| 17| +++++++++++++++++

0| 24| ++++++++++++++++++++++++

250| 16| ++++++++++++++++

500|  4| ++++

750|  2| ++

1000|  0|

1250|  1| +

1500|  0|

1750|  1| +

Histogram - Link Triangle Closure of T PTB NPL ROA , Mean/RMS/SD= -143  568  550

PicoSec|No.|     Total In/Out:  456/   81

-1500|  2| ++

-1250|  0|

-1000|  5| +++++

-750|  5| +++++

-500| 18| ++++++++++++++++++

-250| 16| ++++++++++++++++

0| 13| +++++++++++++

250|  9| +++++++++

500|  5| +++++

750|  3| +++

1000|  4| ++++

1250|  1| +

Histogram - Link Triangle Closure of T PTB NPL SP  , Mean/RMS/SD=  -23  543  542

PicoSec|No.|     Total In/Out:  456/   89

-2000|  1| +

-1750|  0|

-1500|  1| +

-1250|  2| ++

-1000|  3| +++

-750|  5| +++++

-500|  5| +++++

-250| 18| ++++++++++++++++++

0| 16| ++++++++++++++++

250| 19| +++++++++++++++++++

500|  9| +++++++++

750|  6| ++++++

1000|  4| ++++

Histogram - Link Triangle Closure of T PTB OP ROA  , Mean/RMS/SD= -969 1076  469

PicoSec|No.|     Total In/Out:  456/  427

-2750|  1| +

-2500|  1| +

-2250|  3| +++

-2000| 10| ++++++++++

-1750| 21| +++++++++++++++++++++

-1500| 39| +++++++++++++++++++++++++++++++++++++++

-1250| 71| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-1000| 96| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-750| 90| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

-500| 53| +++++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 27| +++++++++++++++++++++++++++

0| 13| +++++++++++++

250|  2| ++

Histogram - Link Triangle Closure of T PTB OP SP   , Mean/RMS/SD=  278  598  529

PicoSec|No.|     Total In/Out:  456/  417

-1000|  4| ++++

-750| 16| ++++++++++++++++

-500| 24| ++++++++++++++++++++++++

-250| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 62| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 75| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 75| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

750| 57| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++

1000| 26| ++++++++++++++++++++++++++

1250| 19| +++++++++++++++++++

1500|  3| +++

1750|  1| +

2000|  0|

2250|  0|

2500|  0|

2750|  1| +

Histogram - Link Triangle Closure of T PTB OP USNO , Mean/RMS/SD=   10  966  966

PicoSec|No.|     Total In/Out:  456/  408

-2250|  3| +++

-2000|  5| +++++

-1750| 10| ++++++++++

-1500| 20| ++++++++++++++++++++

-1250| 14| ++++++++++++++

-1000| 27| +++++++++++++++++++++++++++

-750| 29| +++++++++++++++++++++++++++++

-500| 39| +++++++++++++++++++++++++++++++++++++++

-250| 32| ++++++++++++++++++++++++++++++++

0| 40| ++++++++++++++++++++++++++++++++++++++++

250| 40| ++++++++++++++++++++++++++++++++++++++++

500| 37| +++++++++++++++++++++++++++++++++++++

750| 31| +++++++++++++++++++++++++++++++

1000| 29| +++++++++++++++++++++++++++++

1250| 21| +++++++++++++++++++++

1500| 11| +++++++++++

1750| 14| ++++++++++++++

2000|  1| +

2250|  2| ++

2500|  1| +

2750|  1| +

3000|  1| +

Histogram - Link Triangle Closure of T PTB OP VSL  , Mean/RMS/SD= -628  972  741

PicoSec|No.|     Total In/Out:  456/  367

-3000|  2| ++

-2750|  0|

-2500|  0|

-2250|  4| ++++

-2000|  6| ++++++

-1750| 19| +++++++++++++++++++

-1500| 22| ++++++++++++++++++++++

-1250| 42| ++++++++++++++++++++++++++++++++++++++++++

-1000| 39| +++++++++++++++++++++++++++++++++++++++

-750| 55| +++++++++++++++++++++++++++++++++++++++++++++++++++++++

-500| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

-250| 41| +++++++++++++++++++++++++++++++++++++++++

0| 35| +++++++++++++++++++++++++++++++++++

250| 20| ++++++++++++++++++++

500| 14| ++++++++++++++

750|  4| ++++

1000|  7| +++++++

1250|  3| +++

1500|  1| +

1750|  1| +

2000|  0|

2250|  1| +

Histogram - Link Triangle Closure of T PTB ROA SP  , Mean/RMS/SD=  604  761  463

PicoSec|No.|     Total In/Out:  456/  414

-1000|  1| +

-750|  2| ++

-500|  4| ++++

-250| 17| +++++++++++++++++

0| 38| ++++++++++++++++++++++++++++++++++++++

250| 54| ++++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 85| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

750|109| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

1000| 63| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

1250| 22| ++++++++++++++++++++++

1500| 12| ++++++++++++

1750|  3| +++

2000|  3| +++

2250|  0|

2500|  1| +

Histogram - Link Triangle Closure of T PTB ROA USNO, Mean/RMS/SD=   20 1016 1015

PicoSec|No.|     Total In/Out:  456/  411

-2500|  1| +

-2250|  1| +

-2000|  8| ++++++++

-1750|  7| +++++++

-1500| 13| +++++++++++++

-1250| 18| ++++++++++++++++++

-1000| 23| +++++++++++++++++++++++

-750| 30| ++++++++++++++++++++++++++++++

-500| 47| +++++++++++++++++++++++++++++++++++++++++++++++

-250| 37| +++++++++++++++++++++++++++++++++++++

0| 37| +++++++++++++++++++++++++++++++++++++

250| 42| ++++++++++++++++++++++++++++++++++++++++++

500| 38| ++++++++++++++++++++++++++++++++++++++

750| 28| ++++++++++++++++++++++++++++

1000| 24| ++++++++++++++++++++++++

1250| 16| ++++++++++++++++

1500| 15| +++++++++++++++

1750| 13| +++++++++++++

2000|  3| +++

2250|  7| +++++++

2500|  3| +++

Histogram - Link Triangle Closure of T PTB ROA VSL , Mean/RMS/SD=   -4  575  575

PicoSec|No.|     Total In/Out:  456/  365

-1750|  1| +

-1500|  1| +

-1250|  7| +++++++

-1000| 13| +++++++++++++

-750| 33| +++++++++++++++++++++++++++++++++

-500| 42| ++++++++++++++++++++++++++++++++++++++++++

-250| 50| ++++++++++++++++++++++++++++++++++++++++++++++++++

0| 69| +++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

250| 53| +++++++++++++++++++++++++++++++++++++++++++++++++++++

500| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

750| 24| ++++++++++++++++++++++++

1000| 13| +++++++++++++

1250|  5| +++++

1500|  2| ++

1750|  0|

2000|  0|

2250|  0|

2500|  1| +

Histogram - Link Triangle Closure of T PTB SP USNO , Mean/RMS/SD=   20  872  872

PicoSec|No.|     Total In/Out:  456/  365

-2750|  2| ++

-2500|  2| ++

-2250|  2| ++

-2000|  4| ++++

-1750|  4| ++++

-1500|  9| +++++++++

-1250|  8| ++++++++

-1000| 20| ++++++++++++++++++++

-750| 26| ++++++++++++++++++++++++++

-500| 41| +++++++++++++++++++++++++++++++++++++++++

-250| 39| +++++++++++++++++++++++++++++++++++++++

0| 30| ++++++++++++++++++++++++++++++

250| 48| ++++++++++++++++++++++++++++++++++++++++++++++++

500| 41| +++++++++++++++++++++++++++++++++++++++++

750| 26| ++++++++++++++++++++++++++

1000| 28| ++++++++++++++++++++++++++++

1250| 17| +++++++++++++++++

1500| 10| ++++++++++

1750|  3| +++

2000|  5| +++++

Histogram - Link Triangle Closure of T PTB SP VSL  , Mean/RMS/SD= -472  774  614

PicoSec|No.|     Total In/Out:  456/  352

-2500|  1| +

-2250|  3| +++

-2000|  0|

-1750|  8| ++++++++

-1500| 13| +++++++++++++

-1250| 30| ++++++++++++++++++++++++++++++

-1000| 33| +++++++++++++++++++++++++++++++++

-750| 51| +++++++++++++++++++++++++++++++++++++++++++++++++++

-500| 47| +++++++++++++++++++++++++++++++++++++++++++++++

-250| 62| ++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++++

0| 45| +++++++++++++++++++++++++++++++++++++++++++++

250| 37| +++++++++++++++++++++++++++++++++++++

500| 13| +++++++++++++

750|  4| ++++

1000|  3| +++

1250|  2| ++

Histogram - Link Triangle Closure of T PTB USNO VSL, Mean/RMS/SD=    0 1491 1491

PicoSec|No.|     Total In/Out:  456/  199

-4500|  1| +

-4250|  0|

-4000|  1| +

-3750|  3| +++

-3500|  1| +

-3250|  1| +

-3000|  1| +

-2750|  2| ++

-2500|  4| ++++

-2250|  6| ++++++

-2000|  2| ++

-1750|  4| ++++

-1500|  5| +++++

-1250|  9| +++++++++

-1000| 13| +++++++++++++

-750| 10| ++++++++++

-500| 13| +++++++++++++

-250| 10| ++++++++++

0|  9| +++++++++

250| 17| +++++++++++++++++

500| 17| +++++++++++++++++

750| 16| ++++++++++++++++

1000| 10| ++++++++++

1250| 13| +++++++++++++

1500|  7| +++++++

1750|  5| +++++

2000|  7| +++++++

2250|  1| +

2500|  6| ++++++

2750|  2| ++

3000|  0|

3250|  1| +

3500|  0|

3750|  2| ++

# Text boxes

CI   S   LABi LABj   CALR/ns  U/ns   ESDVAR/ns

171   1   PTB  AOS    165.000   7.0      25.600

1   AOS  PTB   -165.000   7.0   999999999

172   1   PTB  IPQ    188.600   7.0      25.600

1   IPQ  PTB   -188.600   7.0   999999999

173   1   PTB  NPL   -490.300   7.0      25.600

1   NPL  PTB    490.300   7.0   999999999

174   1   PTB USNO   -214.500   3.0      25.600

1  USNO  PTB    214.500   3.0    -379.910

TYPE: G.A. (GPS alignment)

Used Since: MJD 55502 (2 Nov. 2010)

Remark: TM176

CI    Labi Labj   S       CARL   U/ns      ESDVAR  .

175   AOS     CH  S=1      43.722  7.0     999999999

CH    AOS  S=1     -43.722  7.0         8.000

176   AOS     IT  S=1     315.765  7.0     999999999

IT    AOS  S=1    -315.765  7.0     999999999

177   AOS    IPQ  S=1      22.850  7.0     999999999

IPQ    AOS  S=1     -22.850  7.0     999999999

178   AOS   NIST  S=1      31.924  7.0     999999999

NIST    AOS  S=1     -31.924  7.0        -0.724

179   AOS     OP  S=1    7134.602  7.0     999999999

OP    AOS  S=1   -7134.602  7.0     999999999

180   AOS    ROA  S=1     134.121  7.0     999999999

ROA    AOS  S=1    -134.121  7.0       244.4

181   AOS     SP  S=1      29.494  7.0     999999999

SP    AOS  S=1     -29.494  7.0     999999999

182   AOS   USNO  S=1    -380.016  7.0     999999999

USNO    AOS  S=1     380.016  7.0      -379.91

183   AOS    VSL  S=1     306.536  7.0     999999999

VSL    AOS  S=1    -306.536  7.0     999999999

184    CH    IPQ  S=1     -20.772  7.0         8.000

IPQ     CH  S=1      20.772  7.0     999999999

185    CH    NPL  S=1    -698.956  7.0         8.000

NPL     CH  S=1     698.956  7.0     999999999

186    CH   USNO  S=1    -425.517  7.0         8.000

USNO     CH  S=1     425.517  7.0      -379.91

187    IT    IPQ  S=1    -292.613  7.0     999999999

IPQ     IT  S=1     292.613  7.0     999999999

188    IT    NPL  S=1    -971.260  7.0     999999999

NPL     IT  S=1     971.260  7.0     999999999

189    IT   USNO  S=1    -696.084  7.0     999999999

USNO     IT  S=1     696.084  7.0      -379.91

190   IPQ   NIST  S=1       9.179  7.0     999999999

NIST    IPQ  S=1      -9.179  7.0        -0.724

191   IPQ     OP  S=1    7112.281  7.0     999999999

OP    IPQ  S=1   -7112.281  7.0     999999999

192   IPQ    ROA  S=1     110.920  7.0     999999999

ROA    IPQ  S=1    -110.920  7.0       244.4

193   IPQ     SP  S=1       8.266  7.0     999999999

SP    IPQ  S=1      -8.266  7.0     999999999

194   IPQ   USNO  S=1    -404.188  7.0     999999999

USNO    IPQ  S=1     404.188  7.0      -379.91

195   IPQ    VSL  S=1     285.542  7.0     999999999

VSL    IPQ  S=1    -285.542  7.0     999999999

196   NPL     OP  S=1    7790.249  7.0     999999999

OP    NPL  S=1   -7790.249  7.0     999999999

197   NPL    ROA  S=1     788.880  7.0     999999999

ROA    NPL  S=1    -788.880  7.0       244.4

198   NPL     SP  S=1     685.324  7.0     999999999

SP    NPL  S=1    -685.324  7.0     999999999

199    OP   USNO  S=1   -7515.554  7.0     999999999

USNO     OP  S=1    7515.554  7.0      -379.91

200   ROA   USNO  S=1    -516.927  7.0       244.4

USNO    ROA  S=1     516.927  7.0      -379.91

201    SP   USNO  S=1    -411.545  7.0     999999999

USNO     SP  S=1     411.545  7.0      -379.91

202  NIST    ROA  S=1     -18.702  7.0        -0.724

ROA   NIST  S=1      18.702  7.0     999999999

203  USNO    VSL  S=1     686.291  7.0      -379.91

VSL   USNO  S=1    -686.291  7.0     999999999

TYPE: T.C. (triangle closure)

Used Since: MJD 55502 (2 NOV. 2010)

Remark: TM176

Labi Labj        d+/-StDev    N   ESIG

AOS   CH   0.005 +/- 0.726  388 0.037

CH  AOS  -0.005 +/- 0.726  388 0.037

AOS   IT  -0.012 +/- 1.568  372 0.081

IT  AOS   0.012 +/- 1.568  372 0.081

AOS  IPQ  -0.009 +/- 0.795  381 0.041

IPQ  AOS   0.009 +/- 0.795  381 0.041

AOS NIST  -0.001 +/- 0.590  390 0.030

NIST  AOS   0.001 +/- 0.590  390 0.030

AOS   OP   0.006 +/- 1.120  390 0.057

OP  AOS  -0.006 +/- 1.120  390 0.057

AOS  ROA   0.002 +/- 0.723  387 0.037

ROA  AOS  -0.002 +/- 0.723  387 0.037

AOS   SP   0.015 +/- 1.199  375 0.062

SP  AOS  -0.015 +/- 1.199  375 0.062

AOS USNO   0.005 +/- 0.762  379 0.039

USNO  AOS  -0.005 +/- 0.762  379 0.039

AOS  VSL  -0.013 +/- 1.324  328 0.073

VSL  AOS   0.013 +/- 1.324  328 0.073

CH   IT   2.186 +/- 1.530  403 0.076

IT   CH  -2.186 +/- 1.530  403 0.076

CH  IPQ  -0.015 +/- 1.056  419 0.052

IPQ   CH   0.015 +/- 1.056  419 0.052

CH NIST  -0.449 +/- 0.684  431 0.033

NIST   CH   0.449 +/- 0.684  431 0.033

CH  NPL  -0.027 +/- 0.950  118 0.087

NPL   CH   0.027 +/- 0.950  118 0.087

CH   OP   0.150 +/- 0.815  430 0.039

OP   CH  -0.150 +/- 0.815  430 0.039

CH  ROA  -0.685 +/- 0.739  427 0.036

ROA   CH   0.685 +/- 0.739  427 0.036

CH   SP   0.285 +/- 1.140  372 0.059

SP   CH  -0.285 +/- 1.140  372 0.059

CH USNO  -0.004 +/- 0.730  416 0.036

USNO   CH   0.004 +/- 0.730  416 0.036

CH  VSL  -0.658 +/- 1.085   98 0.110

VSL   CH   0.658 +/- 1.085   98 0.110

IT  IPQ   0.009 +/- 1.526  318 0.086

IPQ   IT  -0.009 +/- 1.526  318 0.086

IT NIST  -3.131 +/- 1.098  403 0.055

NIST   IT   3.131 +/- 1.098  403 0.055

IT  NPL  -0.011 +/- 0.573  230 0.038

NPL   IT   0.011 +/- 0.573  230 0.038

IT   OP  -1.120 +/- 0.900  403 0.045

OP   IT   1.120 +/- 0.900  403 0.045

IT  ROA  -0.478 +/- 0.296  400 0.015

ROA   IT   0.478 +/- 0.296  400 0.015

IT   SP  -0.740 +/- 0.849  397 0.043

SP   IT   0.740 +/- 0.849  397 0.043

IT USNO   0.013 +/- 1.732  391 0.088

USNO   IT  -0.013 +/- 1.732  391 0.088

IT  VSL  -0.544 +/- 0.288   10 0.091

VSL   IT   0.544 +/- 0.288   10 0.091

IPQ NIST   0.005 +/- 0.726  419 0.035

NIST  IPQ  -0.005 +/- 0.726  419 0.035

IPQ   OP  -0.002 +/- 1.194  424 0.058

OP  IPQ   0.002 +/- 1.194  424 0.058

IPQ  ROA   0.017 +/- 0.853  420 0.042

ROA  IPQ  -0.017 +/- 0.853  420 0.042

IPQ   SP   0.013 +/- 3.763  271 0.229

SP  IPQ  -0.013 +/- 3.763  271 0.229

IPQ USNO   0.009 +/- 0.953  393 0.048

USNO  IPQ  -0.009 +/- 0.953  393 0.048

IPQ  VSL  -0.137 +/- 9.106  112 0.860

VSL  IPQ   0.137 +/- 9.106  112 0.860

NIST   OP   0.424 +/- 0.693  421 0.034

OP NIST  -0.424 +/- 0.693  421 0.034

NIST  ROA   0.000 +/- 0.661  427 0.032

ROA NIST   0.000 +/- 0.661  427 0.032

NIST   SP   1.733 +/- 0.817  424 0.040

SP NIST  -1.733 +/- 0.817  424 0.040

NIST  VSL   0.316 +/- 1.202  369 0.063

Links[ns with (+) for Link1]: _1007_IPQPTB.TMTA5 , Total: 131Ep

DifLink: Min,Max,Mean: 182.841__194.054__188.584

Sigma= 2.196ns-MeanRmved